## Calculator
U/L2050/60 X-Series USB/LAN Peak and Average Power Sensor
Uncertainty calculation for a power measurement (settled, average power)
Instruction: Fill up blue colour columns
1 | Power Level (dBm) | = | 0 |  | = | 0.001 | W
2 | Frequency (GHz) | = | 6
| Average setting | = | 1
| VBW setting | = | Off
| Free run / Trigger mode | = | Free run mode
| Sensor mode | = | Normal mode | (Normal mode enables peak/average power measurements and trace view)
| Sensor model | = | L2066XA
| DUT SWR | = | 1.26
| Gate length (us) - not required for Free Run or Average-only mode | = | 500
Fields applicable to U/L2060 X-Series only:
| (a) Aperture time (ms) | = | 50
| (b) Internal or External Zero | = | External Zero | (Internal zero - allow zeroing-on-the-go without switching off RF input power) (For U/L2050/60 X-Series, use External Zero when measure signal below -50dBm)
|  | Symbol | Value | Percentage
3 | Calculate sensor uncertainty due to measurement noise and zero drift
| Measurement noise | N | 3.2e-08 | 0.0032%
| Zero drift | D | 1.5e-08 | 0.0015%
| RSS of above terms | RSS(N,D) | -- | 0.003534%
4 | Zero uncertainty | Zs | 2.7e-08 | 0.0027%
5 | Sensor calibration uncertainty | Kb | 4.2% | 4.2%
6 | System contribution, coverage factor of 2 | RSS(N, D, Zs, Kb) | -- | 4.2%
| (RSS of item 3, 4 and 5) | -- | --
7 | Standard uncertainty of mismatch | -- | --
| Sensor max SWR | ¦Гsensor¦ | 1.22
| DUT max SWR | ¦ГDUT¦ | 1.26
8 | Combined measurement uncertainty @ k = 1 | Uc | = | 2.249%
| k factor | k | = | 2
| Expanded Uncertainty | Uc * k | = | 4.499%
|  | Upper Limit Uncertainty | = | 0.19111
|  | Lower Limit Uncertainty | = | 0.19991
Revision 3
Date updated: 12 Sept 2022
© Keysight Technologies, Inc. 2012, 2014, 2015
The data used in this calculation is base on the specifications as in the U/L2050/60 X-Series USB/LAN power sensors datasheet.
This calculation is base on ISO Guide to the Expression of Uncertainty in Measurement, often referred to as the GUM.
For more info on power meter and power sensor measurement uncertainty equation and calculation base on GUM, please refer to Keysight Application Note 1449-3 "Fundamental of RF and Microwave Power Measurements (Part 3), literature number 5988-9215EN.

## Data (hidden)
Measurement noise | Sensor mode
Free run mode | Normal mode
Trigger mode | Average only mode | [7 empty] | U2040 & U2065XA Normal mode
[10 empty] | Low/Med | High/off
Normal mode | [8 empty] | Meas noise | 1e-08 | 3.2e-08
Free run (U2040) | [8 empty] | Noise per sample | 1.5e-07 | 8e-07
Noise | 1e-07
Noise Multiplier | 1 | 2 | 4 | 8 | 16 | 32 | 64 | 128 | 256 | 512 | 1024
| 1 | 0.9 | 0.8 | 0.7 | 0.6 | 0.5 | 0.45 | 0.4 | 0.3 | 0.25 | 0.2
Free run measurement noise | = measurement noise x free run noise multiplier
(U2020) | 1e-07 | W | [6 empty] | Free run meas noise |  | =measurement noise x free run noise multiplier
Trigger mode | [8 empty] | (U2040) & (U2065XA) | = | 3.2e-08 | W
Noise per sample | 0.05 | 0.499 | 0.5 | 0.501 | 40
(U2020) | 3e-06 | 3e-06 | 2.5e-06 | 2.5e-06 | 2.5e-06
Noise per sample multiplier |  | <500MHz | >=500MHz
| Low | 0.6 | 0.55
| Med | 1.3 | 0.65
| High | 2.7 | 0.8
| Off | 1 | 1
Trigger mode measurement noise | = noise per sample x noise per sample multiplier x 1/sqrt[(gate length/12.5ns)] | [7 empty] | Trigger mode meas noise |  |  | = noise per sample x noise per saple multiplier x 1/sqrt[gate length/50ns]
(U2020) | 2.5e-06 | x | 1 | x | 0.005 |  |  |  | (U2040) & (U2065XA) |  | = | 8e-07 | x | 0.01
= | 1.25e-08 | W | [8 empty] | = | 8e-09 | W
Trigger mode measurement noise improvement limits at the measurement noise of 100 nW:
= | 1e-07
U2020 NORM Measurement noise = | 1e-07 | W | [6 empty] | U2040 & U2065XA NORM meas noise |  | = | 3.2e-08 | W | 5.6569e-08
NORM measurement noise = | 3.2e-08 | W
U2020 Final Measurement noise in % | = Noise / Power, noise power is capped for power > 100 uW, for these cases, use: Noise / 100 uW
= | 0.1% | [6 empty] | U2040 & U2065XA final NORM mode meas noise in % = |  |  |  | Noise / power
[11 empty] | = | 0.0032%
Final NORM meas noise in % | 0.0032%
NORM mode zero drift | U2021XA | U2022XA | U2051XA | U2052XA | U2053XA | U2054XA | U2055XA | U2056XA | U2057XA | U2061XA | U2062XA | U2063XA | U2064XA | U2065XA | U2066XA | U2067XA | L2051XA | L2052XA | L2053XA | L2054XA | L2055XA | L2056XA | L2057XA | L2061XA | L2062XA | L2063XA | L2064XA | L2065XA | L2066XA | L2067XA | L2065XT | L2066XT | L2067XT
| 1e-07 | 1e-07 | 1.5e-08 | 1.5e-08 | 1.5e-08 | 1.5e-08 | 1.5e-08 | 1.5e-08 | 1.5e-08 | 1.5e-08 | 1.5e-08 | 1.5e-08 | 1.5e-08 | 1.5e-08 | 1.5e-08 | 1.5e-08 | 1.5e-08 | 1.5e-08 | 1.5e-08 | 1.5e-08 | 1.5e-08 | 1.5e-08 | 1.5e-08 | 1.5e-08 | 1.5e-08 | 1.5e-08 | 1.5e-08 | 1.5e-08 | 1.5e-08 | 1.5e-08 | 1.5e-08 | 1.5e-08 | 1.5e-08
NORM mode zero drift = | 0.0032% U2021XA: 1.5e-08; 0.0032% U2022XA: W
[35 empty] | U2040 NORM Zero Set |  |  |  | U/L2050/60XA NORM Zero Set
NORM mode Zero set | U2021XA | U2022XA | U2051XA | U2052XA | U2053XA | U2054XA | U2055XA | U2056XA | U2057XA | U2061XA | U2062XA | U2063XA | U2064XA | U2065XA | U2066XA | U2067XA | L2051XA | L2052XA | L2053XA | L2054XA | L2055XA | L2056XA | L2057XA | L2061XA | L2062XA | L2063XA | L2064XA | L2065XA | L2066XA | L2067XA | L2065XT | L2066XT | L2067XT |  |  | External Zero | Internal Zero |  |  | External Zero | Internal Zero
| 2e-07 | 2e-07 | 2.7e-08 | 2.7e-08 | 2.7e-08 | 2.7e-08 | 2.7e-08 | 2.7e-08 | 2.7e-08 | 2.7e-08 | 2.7e-08 | 2.7e-08 | 2.7e-08 | 2.7e-08 | 2.7e-08 | 2.7e-08 | 2.7e-08 | 2.7e-08 | 2.7e-08 | 2.7e-08 | 2.7e-08 | 2.7e-08 | 2.7e-08 | 2.7e-08 | 2.7e-08 | 2.7e-08 | 2.7e-08 | 2.7e-08 | 2.7e-08 | 2.7e-08 | 2.7e-08 | 2.7e-08 | 2.7e-08 |  | Low | 1.6e-08 | 2.3e-08 |  | Low | 1.2e-08 | 1.5e-08
NORM mode zero set = | U2021XA: 2.7e-08; U2022XA: W |  | Med | 1.6e-08 | 2.3e-08 |  | Med | 1.2e-08 | 1.5e-08
[35 empty] | High | 5e-08 | 6e-08 |  | High | 2.7e-08 | 3e-08
[35 empty] | Off | 5e-08 | 6e-08 |  | Off | 2.7e-08 | 3e-08
Calibration uncertainty
| U2021XA | U2022XA
0.05 | U2021XA: 4.2%; U2022XA: 4.3%
0.5 | U2021XA: 4.2%; U2022XA: 4.3%
0.5001 | U2021XA: 4%; U2022XA: 4.2%
1.001 | U2021XA: 4%; U2022XA: 4.5%
10.001 | U2021XA: 4.5%; U2022XA: 4.5%
18.001 | U2021XA: #; U2022XA: 5.3%; U2052XA: U2020 Cal Factor =; U2054XA: 4.5%
26.501 | U2021XA: #; U2022XA: 5.8%
40 | U2021XA: #; U2022XA: 5.8%
| L2067XA: U/L2054/55/64X/65 (-70 to +10dBm) | [20 empty] | U/L2054/55/64X/65 (-70 to +10dBm)
Sensor max SWR | U2054XA: BW58:AO58; U2055XA: BX58:AP58; L2064XA: U/L2051/52/53/61/62/63 (-70 to +15dBm) | [17 empty] | U/L2051/52/53/61/62/63 (-70 to +15dBm)
| U2021XA | U2022XA | U2051XA | U2052XA | U2053XA | U2054XA | U2055XA | U2056XA | U2057XA | U2061XA | U2062XA | U2063XA | U2064XA | U2065XA | U2066XA | U2067XA | L2051XA | L2052XA | L2053XA | L2054XA | L2055XA | L2056XA | L2057XA | L2061XA | L2062XA | L2063XA | L2064XA | L2065XA | L2066XA | L2067XA | L2065XT | L2066XT | L2067XT |  |  |  | U2051XA | U2052XA | U2053XA | U2054XA | U2055XA | U2056XA | U2057XA | U2061XA | U2062XA | U2063XA | U2064XA | U2065XA | U2066XA | U2067XA | L2051XA | L2052XA | L2053XA | L2054XA | L2055XA | L2056XA | L2057XA | L2061XA | L2062XA | L2063XA | L2065XA | L2065XA | L2066XA | L2067XA | L2065XT | L2066XT | L2067XT |  |  |  | U2051XA | U2052XA | U2053XA | U2054XA | U2055XA | U2056XA | U2057XA | U2061XA | U2062XA | U2063XA | U2064XA | U2065XA | U2066XA | U2067XA | L2051XA | L2052XA | L2053XA | L2054XA | L2055XA | L2056XA | L2057XA | L2061XA | L2062XA | L2063XA | L2065XA | L2065XA | L2066XA | L2067XA | L2065XT | L2066XT | L2067XT
0.01 | # | # | 1.15 | 1.15 | 1.16 | 1.6 | 1.6 | 2.22 | 2.22 | 1.15 | 1.15 | 1.16 | 1.6 | 1.6 | 2.22 | 2.22 | 1.15 | 1.15 | 1.16 | 1.6 | 1.6 | 2.22 | 2.22 | 1.15 | 1.15 | 1.16 | 1.6 | 1.6 | 2.22 | 2.22 | 1.6 | 2.22 | 2.22 |  |  | 0.01 | 1.15 | 1.15 | 1.16 | 1.6 | 1.6 | 2.22 | 2.22 | 1.15 | 1.15 | 1.16 | 1.6 | 1.6 | 2.22 | 2.22 | 1.15 | 1.15 | 1.16 | 1.6 | 1.6 | 2.22 | 2.22 | 1.15 | 1.15 | 1.16 | 1.6 | 1.6 | 2.22 | 2.22 | 1.6 | 2.22 | 2.22 |  |  | 0.01 | 1.24 | 1.24 | 1.24 | 1.6 | 1.6 | 2.22 | 2.22 | 1.24 | 1.24 | 1.24 | 1.6 | 1.6 | 2.22 | 2.22 | 1.24 | 1.24 | 1.24 | 1.6 | 1.6 | 2.22 | 2.22 | 1.24 | 1.24 | 1.24 | 1.6 | 1.6 | 2.22 | 2.22 | 1.6 | 2.22 | 2.22
0.0301 | # | # | 1.15 | 1.15 | 1.16 | 1.15 | 1.15 | 1.24 | 1.24 | 1.15 | 1.15 | 1.16 | 1.15 | 1.15 | 1.24 | 1.24 | 1.15 | 1.15 | 1.16 | 1.15 | 1.15 | 1.24 | 1.24 | 1.15 | 1.15 | 1.16 | 1.15 | 1.15 | 1.24 | 1.24 | 1.15 | 1.24 | 1.24 |  |  | 0.0301 | 1.15 | 1.15 | 1.16 | 1.15 | 1.15 | 1.24 | 1.24 | 1.15 | 1.15 | 1.16 | 1.15 | 1.15 | 1.24 | 1.24 | 1.15 | 1.15 | 1.16 | 1.15 | 1.15 | 1.24 | 1.24 | 1.15 | 1.15 | 1.16 | 1.15 | 1.15 | 1.24 | 1.24 | 1.15 | 1.24 | 1.24 |  |  | 0.0301 | 1.24 | 1.24 | 1.24 | 1.22 | 1.22 | 1.24 | 1.24 | 1.24 | 1.24 | 1.24 | 1.22 | 1.22 | 1.24 | 1.24 | 1.24 | 1.24 | 1.24 | 1.22 | 1.22 | 1.24 | 1.24 | 1.24 | 1.24 | 1.24 | 1.22 | 1.22 | 1.24 | 1.24 | 1.22 | 1.24 | 1.24
0.05 | 1.2 | 1.2 | 1.15 | 1.15 | 1.16 | 1.13 | 1.15 | 1.24 | 1.24 | 1.15 | 1.15 | 1.16 | 1.13 | 1.15 | 1.24 | 1.24 | 1.15 | 1.15 | 1.16 | 1.13 | 1.15 | 1.24 | 1.24 | 1.15 | 1.15 | 1.16 | 1.13 | 1.15 | 1.24 | 1.24 | 1.15 | 1.24 | 1.24 |  |  | 0.05 | 1.15 | 1.15 | 1.16 | 1.13 | 1.15 | 1.24 | 1.24 | 1.15 | 1.15 | 1.16 | 1.13 | 1.15 | 1.24 | 1.24 | 1.15 | 1.15 | 1.16 | 1.13 | 1.15 | 1.24 | 1.24 | 1.15 | 1.15 | 1.16 | 1.13 | 1.15 | 1.24 | 1.24 | 1.15 | 1.24 | 1.24 |  |  | 0.05 | 1.24 | 1.24 | 1.24 | 1.22 | 1.22 | 1.24 | 1.24 | 1.24 | 1.24 | 1.24 | 1.22 | 1.22 | 1.24 | 1.24 | 1.24 | 1.24 | 1.24 | 1.22 | 1.22 | 1.24 | 1.24 | 1.24 | 1.24 | 1.24 | 1.22 | 1.22 | 1.24 | 1.24 | 1.22 | 1.24 | 1.24
0.075 |  |  | 1.15 | 1.15 | 1.16 | 1.13 | 1.15 | 1.24 | 1.24 | 1.15 | 1.15 | 1.16 | 1.13 | 1.15 | 1.24 | 1.24 | 1.15 | 1.15 | 1.16 | 1.13 | 1.15 | 1.24 | 1.24 | 1.15 | 1.15 | 1.16 | 1.13 | 1.15 | 1.24 | 1.24 | 1.15 | 1.24 | 1.24 |  |  | 0.075 | 1.15 | 1.15 | 1.16 | 1.13 | 1.15 | 1.24 | 1.24 | 1.15 | 1.15 | 1.16 | 1.13 | 1.15 | 1.24 | 1.24 | 1.15 | 1.15 | 1.16 | 1.13 | 1.15 | 1.24 | 1.24 | 1.15 | 1.15 | 1.16 | 1.13 | 1.15 | 1.24 | 1.24 | 1.15 | 1.24 | 1.24 |  |  | 0.075 | 1.24 | 1.24 | 1.24 | 1.22 | 1.22 | 1.24 | 1.24 | 1.24 | 1.24 | 1.24 | 1.22 | 1.22 | 1.24 | 1.24 | 1.24 | 1.24 | 1.24 | 1.22 | 1.22 | 1.24 | 1.24 | 1.24 | 1.24 | 1.24 | 1.22 | 1.22 | 1.24 | 1.24 | 1.22 | 1.24 | 1.24
0.0751 |  |  | 1.15 | 1.15 | 1.16 | 1.13 | 1.15 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.13 | 1.15 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.13 | 1.15 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.13 | 1.15 | 1.19 | 1.19 | 1.15 | 1.19 | 1.19 |  |  | 0.0751 | 1.15 | 1.15 | 1.16 | 1.13 | 1.15 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.13 | 1.15 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.13 | 1.15 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.13 | 1.15 | 1.19 | 1.19 | 1.15 | 1.19 | 1.19 |  |  | 0.0751 | 1.24 | 1.24 | 1.24 | 1.22 | 1.22 | 1.19 | 1.19 | 1.24 | 1.24 | 1.24 | 1.22 | 1.22 | 1.19 | 1.19 | 1.24 | 1.24 | 1.24 | 1.22 | 1.22 | 1.19 | 1.19 | 1.24 | 1.24 | 1.24 | 1.22 | 1.22 | 1.19 | 1.19 | 1.22 | 1.19 | 1.19
0.0501 | 1.2 | 1.2 | 1.15 | 1.15 | 1.16 | 1.13 | 1.13 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.13 | 1.13 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.13 | 1.13 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.13 | 1.13 | 1.19 | 1.19 | 1.13 | 1.19 | 1.19 |  |  | 0.0501 | 1.15 | 1.15 | 1.16 | 1.13 | 1.13 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.13 | 1.13 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.13 | 1.13 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.13 | 1.13 | 1.19 | 1.19 | 1.13 | 1.19 | 1.19 |  |  | 0.0501 | 1.24 | 1.24 | 1.24 | 1.21 | 1.21 | 1.19 | 1.19 | 1.24 | 1.24 | 1.24 | 1.21 | 1.21 | 1.19 | 1.19 | 1.24 | 1.24 | 1.24 | 1.21 | 1.21 | 1.19 | 1.19 | 1.24 | 1.24 | 1.24 | 1.21 | 1.21 | 1.19 | 1.19 | 1.21 | 1.19 | 1.19
0.1 | 1.2 | 1.2 | 1.15 | 1.15 | 1.16 | 1.13 | 1.13 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.13 | 1.13 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.13 | 1.13 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.13 | 1.13 | 1.19 | 1.19 | 1.13 | 1.19 | 1.19 |  |  | 0.1 | 1.15 | 1.15 | 1.16 | 1.13 | 1.13 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.13 | 1.13 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.13 | 1.13 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.13 | 1.13 | 1.19 | 1.19 | 1.13 | 1.19 | 1.19 |  |  | 0.1 | 1.24 | 1.24 | 1.24 | 1.21 | 1.21 | 1.19 | 1.19 | 1.24 | 1.24 | 1.24 | 1.21 | 1.21 | 1.19 | 1.19 | 1.24 | 1.24 | 1.24 | 1.21 | 1.21 | 1.19 | 1.19 | 1.24 | 1.24 | 1.24 | 1.21 | 1.21 | 1.19 | 1.19 | 1.21 | 1.19 | 1.19
0.101 | 1.2 | 1.2 | 1.15 | 1.15 | 1.16 | 1.13 | 1.13 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.13 | 1.13 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.13 | 1.13 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.13 | 1.13 | 1.19 | 1.19 | 1.13 | 1.19 | 1.19 |  |  | 0.101 | 1.15 | 1.15 | 1.16 | 1.13 | 1.13 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.13 | 1.13 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.13 | 1.13 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.13 | 1.13 | 1.19 | 1.19 | 1.13 | 1.19 | 1.19 |  |  | 0.101 | 1.24 | 1.24 | 1.24 | 1.21 | 1.21 | 1.19 | 1.19 | 1.24 | 1.24 | 1.24 | 1.21 | 1.21 | 1.19 | 1.19 | 1.24 | 1.24 | 1.24 | 1.21 | 1.21 | 1.19 | 1.19 | 1.24 | 1.24 | 1.24 | 1.21 | 1.21 | 1.19 | 1.19 | 1.21 | 1.19 | 1.19
0.3 | 1.2 | 1.2 | 1.15 | 1.15 | 1.16 | 1.13 | 1.13 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.13 | 1.13 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.13 | 1.13 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.13 | 1.13 | 1.19 | 1.19 | 1.13 | 1.19 | 1.19 |  |  | 0.3 | 1.15 | 1.15 | 1.16 | 1.13 | 1.13 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.13 | 1.13 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.13 | 1.13 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.13 | 1.13 | 1.19 | 1.19 | 1.13 | 1.19 | 1.19 |  |  | 0.3 | 1.24 | 1.24 | 1.24 | 1.21 | 1.21 | 1.19 | 1.19 | 1.24 | 1.24 | 1.24 | 1.21 | 1.21 | 1.19 | 1.19 | 1.24 | 1.24 | 1.24 | 1.21 | 1.21 | 1.19 | 1.19 | 1.24 | 1.24 | 1.24 | 1.21 | 1.21 | 1.19 | 1.19 | 1.21 | 1.19 | 1.19
0.301 | 1.2 | 1.2 | 1.15 | 1.15 | 1.16 | 1.17 | 1.14 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.17 | 1.14 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.17 | 1.14 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.17 | 1.14 | 1.19 | 1.19 | 1.14 | 1.19 | 1.19 |  |  | 0.301 | 1.15 | 1.15 | 1.16 | 1.17 | 1.14 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.17 | 1.14 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.17 | 1.14 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.17 | 1.14 | 1.19 | 1.19 | 1.14 | 1.19 | 1.19 |  |  | 0.301 | 1.24 | 1.24 | 1.24 | 1.26 | 1.2 | 1.19 | 1.19 | 1.24 | 1.24 | 1.24 | 1.26 | 1.2 | 1.19 | 1.19 | 1.24 | 1.24 | 1.24 | 1.26 | 1.2 | 1.19 | 1.19 | 1.24 | 1.24 | 1.24 | 1.26 | 1.2 | 1.19 | 1.19 | 1.2 | 1.19 | 1.19
2 |  |  | 1.15 | 1.15 | 1.16 | 1.17 | 1.14 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.17 | 1.14 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.17 | 1.14 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.17 | 1.14 | 1.19 | 1.19 | 1.14 | 1.19 | 1.19 |  |  | 2 | 1.15 | 1.15 | 1.16 | 1.17 | 1.14 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.17 | 1.14 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.17 | 1.14 | 1.19 | 1.19 | 1.15 | 1.15 | 1.16 | 1.17 | 1.14 | 1.19 | 1.19 | 1.14 | 1.19 | 1.19 |  |  | 2 | 1.24 | 1.24 | 1.24 | 1.26 | 1.2 | 1.19 | 1.19 | 1.24 | 1.24 | 1.24 | 1.26 | 1.2 | 1.19 | 1.19 | 1.24 | 1.24 | 1.24 | 1.26 | 1.2 | 1.19 | 1.19 | 1.24 | 1.24 | 1.24 | 1.26 | 1.2 | 1.19 | 1.19 | 1.2 | 1.19 | 1.19
2.001 |  |  | 1.15 | 1.15 | 1.16 | 1.17 | 1.14 | 1.22 | 1.22 | 1.15 | 1.15 | 1.16 | 1.17 | 1.14 | 1.22 | 1.22 | 1.15 | 1.15 | 1.16 | 1.17 | 1.14 | 1.22 | 1.22 | 1.15 | 1.15 | 1.16 | 1.17 | 1.14 | 1.22 | 1.22 | 1.14 | 1.22 | 1.22 |  |  | 2.001 | 1.15 | 1.15 | 1.16 | 1.17 | 1.14 | 1.22 | 1.22 | 1.15 | 1.15 | 1.16 | 1.17 | 1.14 | 1.22 | 1.22 | 1.15 | 1.15 | 1.16 | 1.17 | 1.14 | 1.22 | 1.22 | 1.15 | 1.15 | 1.16 | 1.17 | 1.14 | 1.22 | 1.22 | 1.14 | 1.22 | 1.22 |  |  | 2.001 | 1.24 | 1.24 | 1.24 | 1.26 | 1.2 | 1.22 | 1.22 | 1.24 | 1.24 | 1.24 | 1.26 | 1.2 | 1.22 | 1.22 | 1.24 | 1.24 | 1.24 | 1.26 | 1.2 | 1.22 | 1.22 | 1.24 | 1.24 | 1.24 | 1.26 | 1.2 | 1.22 | 1.22 | 1.2 | 1.22 | 1.22
4 | 1.2 | 1.2 | 1.15 | 1.15 | 1.16 | 1.17 | 1.14 | 1.22 | 1.22 | 1.15 | 1.15 | 1.16 | 1.17 | 1.14 | 1.22 | 1.22 | 1.15 | 1.15 | 1.16 | 1.17 | 1.14 | 1.22 | 1.22 | 1.15 | 1.15 | 1.16 | 1.17 | 1.14 | 1.22 | 1.22 | 1.14 | 1.22 | 1.22 |  |  | 4 | 1.15 | 1.15 | 1.16 | 1.17 | 1.14 | 1.22 | 1.22 | 1.15 | 1.15 | 1.16 | 1.17 | 1.14 | 1.22 | 1.22 | 1.15 | 1.15 | 1.16 | 1.17 | 1.14 | 1.22 | 1.22 | 1.15 | 1.15 | 1.16 | 1.17 | 1.14 | 1.22 | 1.22 | 1.14 | 1.22 | 1.22 |  |  | 4 | 1.24 | 1.24 | 1.24 | 1.26 | 1.2 | 1.22 | 1.22 | 1.24 | 1.24 | 1.24 | 1.26 | 1.2 | 1.22 | 1.22 | 1.24 | 1.24 | 1.24 | 1.26 | 1.2 | 1.22 | 1.22 | 1.24 | 1.24 | 1.24 | 1.26 | 1.2 | 1.22 | 1.22 | 1.2 | 1.22 | 1.22
4.001 | 1.2 | 1.2 | 1.15 | 1.15 | 1.16 | 1.21 | 1.16 | 1.22 | 1.22 | 1.15 | 1.15 | 1.16 | 1.21 | 1.16 | 1.22 | 1.22 | 1.15 | 1.15 | 1.16 | 1.21 | 1.16 | 1.22 | 1.22 | 1.15 | 1.15 | 1.16 | 1.21 | 1.16 | 1.22 | 1.22 | 1.16 | 1.22 | 1.22 |  |  | 4.001 | 1.15 | 1.15 | 1.16 | 1.21 | 1.16 | 1.22 | 1.22 | 1.15 | 1.15 | 1.16 | 1.21 | 1.16 | 1.22 | 1.22 | 1.15 | 1.15 | 1.16 | 1.21 | 1.16 | 1.22 | 1.22 | 1.15 | 1.15 | 1.16 | 1.21 | 1.16 | 1.22 | 1.22 | 1.16 | 1.22 | 1.22 |  |  | 4.001 | 1.24 | 1.24 | 1.24 | 1.22 | 1.2 | 1.22 | 1.22 | 1.24 | 1.24 | 1.24 | 1.22 | 1.2 | 1.22 | 1.22 | 1.24 | 1.24 | 1.24 | 1.22 | 1.2 | 1.22 | 1.22 | 1.24 | 1.24 | 1.24 | 1.22 | 1.2 | 1.22 | 1.22 | 1.2 | 1.22 | 1.22
6 | 1.2 | 1.2 | 1.15 | 1.15 | 1.16 | 1.21 | 1.16 | 1.22 | 1.22 | 1.15 | 1.15 | 1.16 | 1.21 | 1.16 | 1.22 | 1.22 | 1.15 | 1.15 | 1.16 | 1.21 | 1.16 | 1.22 | 1.22 | 1.15 | 1.15 | 1.16 | 1.21 | 1.16 | 1.22 | 1.22 | 1.16 | 1.22 | 1.22 |  |  | 6 | 1.15 | 1.15 | 1.16 | 1.21 | 1.16 | 1.22 | 1.22 | 1.15 | 1.15 | 1.16 | 1.21 | 1.16 | 1.22 | 1.22 | 1.15 | 1.15 | 1.16 | 1.21 | 1.16 | 1.22 | 1.22 | 1.15 | 1.15 | 1.16 | 1.21 | 1.16 | 1.22 | 1.22 | 1.16 | 1.22 | 1.22 |  |  | 6 | 1.24 | 1.24 | 1.24 | 1.22 | 1.2 | 1.22 | 1.22 | 1.24 | 1.24 | 1.24 | 1.22 | 1.2 | 1.22 | 1.22 | 1.24 | 1.24 | 1.24 | 1.22 | 1.2 | 1.22 | 1.22 | 1.24 | 1.24 | 1.24 | 1.22 | 1.2 | 1.22 | 1.22 | 1.2 | 1.22 | 1.22
6.001 | 1.2 | 1.2 | 0 | 1.26 | 1.24 | 1.21 | 1.16 | 1.22 | 1.22 | 0 | 1.26 | 1.24 | 1.21 | 1.16 | 1.22 | 1.22 | 0 | 1.26 | 1.24 | 1.21 | 1.16 | 1.22 | 1.22 | 0 | 1.26 | 1.24 | 1.21 | 1.16 | 1.22 | 1.22 | 1.16 | 1.22 | 1.22 |  |  | 6.001 |  | 1.26 | 1.24 | 1.21 | 1.16 | 1.22 | 1.22 |  | 1.26 | 1.24 | 1.21 | 1.16 | 1.22 | 1.22 |  | 1.26 | 1.24 | 1.21 | 1.16 | 1.22 | 1.22 |  | 1.26 | 1.24 | 1.21 | 1.16 | 1.22 | 1.22 | 1.16 | 1.22 | 1.22 |  |  | 6.001 |  | 1.3 | 1.27 | 1.22 | 1.2 | 1.22 | 1.22 |  | 1.3 | 1.27 | 1.22 | 1.2 | 1.22 | 1.22 |  | 1.3 | 1.27 | 1.22 | 1.2 | 1.22 | 1.22 |  | 1.3 | 1.27 | 1.22 | 1.2 | 1.22 | 1.22 | 1.2 | 1.22 | 1.22
8 | 1.2 | 1.2 | 0 | 1.26 | 1.24 | 1.21 | 1.16 | 1.22 | 1.22 | 0 | 1.26 | 1.24 | 1.21 | 1.16 | 1.22 | 1.22 | 0 | 1.26 | 1.24 | 1.21 | 1.16 | 1.22 | 1.22 | 0 | 1.26 | 1.24 | 1.21 | 1.16 | 1.22 | 1.22 | 1.16 | 1.22 | 1.22 |  |  | 8 |  | 1.26 | 1.24 | 1.21 | 1.16 | 1.22 | 1.22 |  | 1.26 | 1.24 | 1.21 | 1.16 | 1.22 | 1.22 |  | 1.26 | 1.24 | 1.21 | 1.16 | 1.22 | 1.22 |  | 1.26 | 1.24 | 1.21 | 1.16 | 1.22 | 1.22 | 1.16 | 1.22 | 1.22 |  |  | 8 |  | 1.3 | 1.27 | 1.22 | 1.2 | 1.22 | 1.22 |  | 1.3 | 1.27 | 1.22 | 1.2 | 1.22 | 1.22 |  | 1.3 | 1.27 | 1.22 | 1.2 | 1.22 | 1.22 |  | 1.3 | 1.27 | 1.22 | 1.2 | 1.22 | 1.22 | 1.2 | 1.22 | 1.22
8.001 | 1.2 | 1.2 | 0 | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 | 0 | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 | 0 | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 | 0 | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 | 1.2 | 1.22 | 1.22 |  |  | 8.001 |  | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 |  | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 |  | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 |  | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 | 1.2 | 1.22 | 1.22 |  |  | 8.001 |  | 1.3 | 1.27 | 1.25 | 1.21 | 1.22 | 1.22 |  | 1.3 | 1.27 | 1.25 | 1.21 | 1.22 | 1.22 |  | 1.3 | 1.27 | 1.25 | 1.21 | 1.22 | 1.22 |  | 1.3 | 1.27 | 1.25 | 1.21 | 1.22 | 1.22 | 1.21 | 1.22 | 1.22
10 | 1.2 | 1.2 | 0 | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 | 0 | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 | 0 | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 | 0 | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 | 1.2 | 1.22 | 1.22 |  |  | 10 |  | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 |  | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 |  | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 |  | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 | 1.2 | 1.22 | 1.22 |  |  | 10 |  | 1.3 | 1.27 | 1.25 | 1.21 | 1.22 | 1.22 |  | 1.3 | 1.27 | 1.25 | 1.21 | 1.22 | 1.22 |  | 1.3 | 1.27 | 1.25 | 1.21 | 1.22 | 1.22 |  | 1.3 | 1.27 | 1.25 | 1.21 | 1.22 | 1.22 | 1.21 | 1.22 | 1.22
10.001 | 1.26 | 1.26 | 0 | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 | 0 | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 | 0 | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 | 0 | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 | 1.2 | 1.22 | 1.22 |  |  | 10.001 |  | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 |  | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 |  | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 |  | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 | 1.2 | 1.22 | 1.22 |  |  | 10.001 |  | 1.3 | 1.27 | 1.25 | 1.21 | 1.22 | 1.22 |  | 1.3 | 1.27 | 1.25 | 1.21 | 1.22 | 1.22 |  | 1.3 | 1.27 | 1.25 | 1.21 | 1.22 | 1.22 |  | 1.3 | 1.27 | 1.25 | 1.21 | 1.22 | 1.22 | 1.21 | 1.22 | 1.22
11.5 | 1.26 | 1.26 | 0 | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 | 0 | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 | 0 | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 | 0 | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 | 1.2 | 1.22 | 1.22 |  |  | 11.5 |  | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 |  | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 |  | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 |  | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 | 1.2 | 1.22 | 1.22 |  |  | 11.5 |  | 1.3 | 1.27 | 1.25 | 1.21 | 1.22 | 1.22 |  | 1.3 | 1.27 | 1.25 | 1.21 | 1.22 | 1.22 |  | 1.3 | 1.27 | 1.25 | 1.21 | 1.22 | 1.22 |  | 1.3 | 1.27 | 1.25 | 1.21 | 1.22 | 1.22 | 1.21 | 1.22 | 1.22
11.501 | 1.26 | 1.26 | 0 | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 | 0 | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 | 0 | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 | 0 | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 | 1.2 | 1.22 | 1.22 |  |  | 11.501 |  | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 |  | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 |  | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 |  | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 | 1.2 | 1.22 | 1.22 |  |  | 11.501 |  | 1.3 | 1.27 | 1.25 | 1.21 | 1.22 | 1.22 |  | 1.3 | 1.27 | 1.25 | 1.21 | 1.22 | 1.22 |  | 1.3 | 1.27 | 1.25 | 1.21 | 1.22 | 1.22 |  | 1.3 | 1.27 | 1.25 | 1.21 | 1.22 | 1.22 | 1.21 | 1.22 | 1.22
14 | 1.26 | 1.26 | 0 | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 | 0 | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 | 0 | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 | 0 | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 | 1.2 | 1.22 | 1.22 |  |  | 14 |  | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 |  | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 |  | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 |  | 1.26 | 1.24 | 1.19 | 1.2 | 1.22 | 1.22 | 1.2 | 1.22 | 1.22 |  |  | 14 |  | 1.3 | 1.27 | 1.25 | 1.21 | 1.22 | 1.22 |  | 1.3 | 1.27 | 1.25 | 1.21 | 1.22 | 1.22 |  | 1.3 | 1.27 | 1.25 | 1.21 | 1.22 | 1.22 |  | 1.3 | 1.27 | 1.25 | 1.21 | 1.22 | 1.22 | 1.21 | 1.22 | 1.22
14.001 | 1.26 | 1.26 | 0 | 1.26 | 1.24 | 1.28 | 1.29 | 1.29 | 1.29 | 0 | 1.26 | 1.24 | 1.28 | 1.29 | 1.29 | 1.29 | 0 | 1.26 | 1.24 | 1.28 | 1.29 | 1.29 | 1.29 | 0 | 1.26 | 1.24 | 1.28 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29 |  |  | 14.001 |  | 1.26 | 1.24 | 1.28 | 1.29 | 1.29 | 1.29 |  | 1.26 | 1.24 | 1.28 | 1.29 | 1.29 | 1.29 |  | 1.26 | 1.24 | 1.28 | 1.29 | 1.29 | 1.29 |  | 1.26 | 1.24 | 1.28 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29 |  |  | 14.001 |  | 1.3 | 1.27 | 1.31 | 1.29 | 1.29 | 1.29 |  | 1.3 | 1.27 | 1.31 | 1.29 | 1.29 | 1.29 |  | 1.3 | 1.27 | 1.31 | 1.29 | 1.29 | 1.29 |  | 1.3 | 1.27 | 1.31 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29
16 |  |  | 0 | 1.26 | 1.24 | 1.28 | 1.29 | 1.29 | 1.29 | 0 | 1.26 | 1.24 | 1.28 | 1.29 | 1.29 | 1.29 | 0 | 1.26 | 1.24 | 1.28 | 1.29 | 1.29 | 1.29 | 0 | 1.26 | 1.24 | 1.28 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29 |  |  | 16 |  | 1.26 | 1.24 | 1.28 | 1.29 | 1.29 | 1.29 |  | 1.26 | 1.24 | 1.28 | 1.29 | 1.29 | 1.29 |  | 1.26 | 1.24 | 1.28 | 1.29 | 1.29 | 1.29 |  | 1.26 | 1.24 | 1.28 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29 |  |  | 16 |  | 1.3 | 1.27 | 1.31 | 1.29 | 1.29 | 1.29 |  | 1.3 | 1.27 | 1.31 | 1.29 | 1.29 | 1.29 |  | 1.3 | 1.27 | 1.31 | 1.29 | 1.29 | 1.29 |  | 1.3 | 1.27 | 1.31 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29
16.001 |  |  | 0 | 1.26 | 1.33 | 1.28 | 1.29 | 1.29 | 1.29 | 0 | 1.26 | 1.33 | 1.28 | 1.29 | 1.29 | 1.29 | 0 | 1.26 | 1.33 | 1.28 | 1.29 | 1.29 | 1.29 | 0 | 1.26 | 1.33 | 1.28 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29 |  |  | 16.001 |  | 1.26 | 1.33 | 1.28 | 1.29 | 1.29 | 1.29 |  | 1.26 | 1.33 | 1.28 | 1.29 | 1.29 | 1.29 |  | 1.26 | 1.33 | 1.28 | 1.29 | 1.29 | 1.29 |  | 1.26 | 1.33 | 1.28 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29 |  |  | 16.001 |  | 1.3 | 1.4 | 1.31 | 1.29 | 1.29 | 1.29 |  | 1.3 | 1.4 | 1.31 | 1.29 | 1.29 | 1.29 |  | 1.3 | 1.4 | 1.31 | 1.29 | 1.29 | 1.29 |  | 1.3 | 1.4 | 1.31 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29
18 | 1.26 | 1.26 | 0 | 1.26 | 1.33 | 1.28 | 1.29 | 1.29 | 1.29 | 0 | 1.26 | 1.33 | 1.28 | 1.29 | 1.29 | 1.29 | 0 | 1.26 | 1.33 | 1.28 | 1.29 | 1.29 | 1.29 | 0 | 1.26 | 1.33 | 1.28 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29 |  |  | 18 |  | 1.26 | 1.33 | 1.28 | 1.29 | 1.29 | 1.29 |  | 1.26 | 1.33 | 1.28 | 1.29 | 1.29 | 1.29 |  | 1.26 | 1.33 | 1.28 | 1.29 | 1.29 | 1.29 |  | 1.26 | 1.33 | 1.28 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29 |  |  | 18 |  | 1.3 | 1.4 | 1.31 | 1.29 | 1.29 | 1.29 |  | 1.3 | 1.4 | 1.31 | 1.29 | 1.29 | 1.29 |  | 1.3 | 1.4 | 1.31 | 1.29 | 1.29 | 1.29 |  | 1.3 | 1.4 | 1.31 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29
18.001 | # | 1.3 | 0 | 0 | 1.33 | 1.28 | 1.29 | 1.29 | 1.29 | 0 | 0 | 1.33 | 1.28 | 1.29 | 1.29 | 1.29 | 0 | 0 | 1.33 | 1.28 | 1.29 | 1.29 | 1.29 | 0 | 0 | 1.33 | 1.28 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29 |  |  | 18.001 |  |  | 1.33 | 1.28 | 1.29 | 1.29 | 1.29 |  |  | 1.33 | 1.28 | 1.29 | 1.29 | 1.29 |  |  | 1.33 | 1.28 | 1.29 | 1.29 | 1.29 |  |  | 1.33 | 1.28 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29 |  |  | 18.001 |  |  | 1.4 | 1.31 | 1.29 | 1.29 | 1.29 |  |  | 1.4 | 1.31 | 1.29 | 1.29 | 1.29 |  |  | 1.4 | 1.31 | 1.29 | 1.29 | 1.29 |  |  | 1.4 | 1.31 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29
26.5 | # | 1.3 | 0 | 0 | 1.33 | 1.28 | 1.29 | 1.29 | 1.29 | 0 | 0 | 1.33 | 1.28 | 1.29 | 1.29 | 1.29 | 0 | 0 | 1.33 | 1.28 | 1.29 | 1.29 | 1.29 | 0 | 0 | 1.33 | 1.28 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29 |  |  | 26.5 |  |  | 1.33 | 1.28 | 1.29 | 1.29 | 1.29 |  |  | 1.33 | 1.28 | 1.29 | 1.29 | 1.29 |  |  | 1.33 | 1.28 | 1.29 | 1.29 | 1.29 |  |  | 1.33 | 1.28 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29 |  |  | 26.5 |  |  | 1.4 | 1.31 | 1.29 | 1.29 | 1.29 |  |  | 1.4 | 1.31 | 1.29 | 1.29 | 1.29 |  |  | 1.4 | 1.31 | 1.29 | 1.29 | 1.29 |  |  | 1.4 | 1.31 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29 | 1.29
26.501 | # | 1.5 | 0 | 0 | 1.41 | 1.36 | 1.32 | 1.31 | 1.31 | 0 | 0 | 1.41 | 1.36 | 1.32 | 1.31 | 1.31 | 0 | 0 | 1.41 | 1.36 | 1.32 | 1.31 | 1.31 | 0 | 0 | 1.41 | 1.36 | 1.32 | 1.31 | 1.31 | 1.32 | 1.31 | 1.31 |  |  | 26.501 |  |  | 1.41 | 1.36 | 1.32 | 1.31 | 1.31 |  |  | 1.41 | 1.36 | 1.32 | 1.31 | 1.31 |  |  | 1.41 | 1.36 | 1.32 | 1.31 | 1.31 |  |  | 1.41 | 1.36 | 1.32 | 1.31 | 1.31 | 1.32 | 1.31 | 1.31 |  |  | 26.501 |  |  | 1.53 | 1.39 | 1.32 | 1.31 | 1.31 |  |  | 1.53 | 1.39 | 1.32 | 1.31 | 1.31 |  |  | 1.53 | 1.39 | 1.32 | 1.31 | 1.31 |  |  | 1.53 | 1.39 | 1.32 | 1.31 | 1.31 | 1.32 | 1.31 | 1.31
30 | # | 2.5 | 0 | 0 | 1.41 | 1.36 | 1.32 | 1.31 | 1.31 | 0 | 0 | 1.41 | 1.36 | 1.32 | 1.31 | 1.31 | 0 | 0 | 1.41 | 1.36 | 1.32 | 1.31 | 1.31 | 0 | 0 | 1.41 | 1.36 | 1.32 | 1.31 | 1.31 | 1.32 | 1.31 | 1.31 |  |  | 30 |  |  | 1.41 | 1.36 | 1.32 | 1.31 | 1.31 |  |  | 1.41 | 1.36 | 1.32 | 1.31 | 1.31 |  |  | 1.41 | 1.36 | 1.32 | 1.31 | 1.31 |  |  | 1.41 | 1.36 | 1.32 | 1.31 | 1.31 | 1.32 | 1.31 | 1.31 |  |  | 30 |  |  | 1.53 | 1.39 | 1.32 | 1.31 | 1.31 |  |  | 1.53 | 1.39 | 1.32 | 1.31 | 1.31 |  |  | 1.53 | 1.39 | 1.32 | 1.31 | 1.31 |  |  | 1.53 | 1.39 | 1.32 | 1.31 | 1.31 | 1.32 | 1.31 | 1.31
30.001 | # | 3.5 | 0 | 0 | 1.41 | 1.36 | 1.32 | 1.31 | 1.31 | 0 | 0 | 1.41 | 1.36 | 1.32 | 1.31 | 1.31 | 0 | 0 | 1.41 | 1.36 | 1.32 | 1.31 | 1.31 | 0 | 0 | 1.41 | 1.36 | 1.32 | 1.31 | 1.31 | 1.32 | 1.31 | 1.31 |  |  | 30.001 |  |  | 1.41 | 1.36 | 1.32 | 1.31 | 1.31 |  |  | 1.41 | 1.36 | 1.32 | 1.31 | 1.31 |  |  | 1.41 | 1.36 | 1.32 | 1.31 | 1.31 |  |  | 1.41 | 1.36 | 1.32 | 1.31 | 1.31 | 1.32 | 1.31 | 1.31 |  |  | 30.001 |  |  | 1.53 | 1.39 | 1.32 | 1.31 | 1.31 |  |  | 1.53 | 1.39 | 1.32 | 1.31 | 1.31 |  |  | 1.53 | 1.39 | 1.32 | 1.31 | 1.31 |  |  | 1.53 | 1.39 | 1.32 | 1.31 | 1.31 | 1.32 | 1.31 | 1.31
33 | # | 4.5 | 0 | 0 | 1.41 | 1.36 | 1.32 | 1.31 | 1.31 | 0 | 0 | 1.41 | 1.36 | 1.32 | 1.31 | 1.31 | 0 | 0 | 1.41 | 1.36 | 1.32 | 1.31 | 1.31 | 0 | 0 | 1.41 | 1.36 | 1.32 | 1.31 | 1.31 | 1.32 | 1.31 | 1.31 |  |  | 33 |  |  | 1.41 | 1.36 | 1.32 | 1.31 | 1.31 |  |  | 1.41 | 1.36 | 1.32 | 1.31 | 1.31 |  |  | 1.41 | 1.36 | 1.32 | 1.31 | 1.31 |  |  | 1.41 | 1.36 | 1.32 | 1.31 | 1.31 | 1.32 | 1.31 | 1.31 |  |  | 33 |  |  | 1.53 | 1.39 | 1.32 | 1.31 | 1.31 |  |  | 1.53 | 1.39 | 1.32 | 1.31 | 1.31 |  |  | 1.53 | 1.39 | 1.32 | 1.31 | 1.31 |  |  | 1.53 | 1.39 | 1.32 | 1.31 | 1.31 | 1.32 | 1.31 | 1.31
33.001 | # | 5.5 | 0 | 0 | 0 | 1.36 | 1.32 | 1.31 | 1.31 | 0 | 0 | 0 | 1.36 | 1.32 | 1.31 | 1.31 | 0 | 0 | 0 | 1.36 | 1.32 | 1.31 | 1.31 | 0 | 0 | 0 | 1.36 | 1.32 | 1.31 | 1.31 | 1.32 | 1.31 | 1.31 |  |  | 33.001 |  |  |  | 1.36 | 1.32 | 1.31 | 1.31 |  |  |  | 1.36 | 1.32 | 1.31 | 1.31 |  |  |  | 1.36 | 1.32 | 1.31 | 1.31 |  |  |  | 1.36 | 1.32 | 1.31 | 1.31 | 1.32 | 1.31 | 1.31 |  |  | 33.001 |  |  |  | 1.39 | 1.32 | 1.31 | 1.31 |  |  |  | 1.39 | 1.32 | 1.31 | 1.31 |  |  |  | 1.39 | 1.32 | 1.31 | 1.31 |  |  |  | 1.39 | 1.32 | 1.31 | 1.31 | 1.32 | 1.31 | 1.31
40 | # | 6.5 | 0 | 0 | 0 | 1.36 | 1.32 | 1.31 | 1.31 | 0 | 0 | 0 | 1.36 | 1.32 | 1.31 | 1.31 | 0 | 0 | 0 | 1.36 | 1.32 | 1.31 | 1.31 | 0 | 0 | 0 | 1.36 | 1.32 | 1.31 | 1.31 | 1.32 | 1.31 | 1.31 |  |  | 40 |  |  |  | 1.36 | 1.32 | 1.31 | 1.31 |  |  |  | 1.36 | 1.32 | 1.31 | 1.31 |  |  |  | 1.36 | 1.32 | 1.31 | 1.31 |  |  |  | 1.36 | 1.32 | 1.31 | 1.31 | 1.32 | 1.31 | 1.31 |  |  | 40 |  |  |  | 1.39 | 1.32 | 1.31 | 1.31 |  |  |  | 1.39 | 1.32 | 1.31 | 1.31 |  |  |  | 1.39 | 1.32 | 1.31 | 1.31 |  |  |  | 1.39 | 1.32 | 1.31 | 1.31 | 1.32 | 1.31 | 1.31
40.001 | # | # | 0 | 0 | 0 | 0 | 1.4 | 1.33 | 1.33 | 0 | 0 | 0 | 0 | 1.4 | 1.33 | 1.33 | 0 | 0 | 0 | 0 | 1.4 | 1.33 | 1.33 | 0 | 0 | 0 | 0 | 1.4 | 1.33 | 1.33 | 1.4 | 1.33 | 1.33 |  |  | 40.001 |  |  |  |  | 1.4 | 1.33 | 1.33 |  |  |  |  | 1.4 | 1.33 | 1.33 |  |  |  |  | 1.4 | 1.33 | 1.33 |  |  |  |  | 1.4 | 1.33 | 1.33 | 1.4 | 1.33 | 1.33 |  |  | 40.001 |  |  |  |  | 1.4 | 1.33 | 1.33 |  |  |  |  | 1.4 | 1.33 | 1.33 |  |  |  |  | 1.4 | 1.33 | 1.33 |  |  |  |  | 1.4 | 1.33 | 1.33 | 1.4 | 1.33 | 1.33
48 | # | # | 0 | 0 | 0 | 0 | 1.4 | 1.33 | 1.33 | 0 | 0 | 0 | 0 | 1.4 | 1.33 | 1.33 | 0 | 0 | 0 | 0 | 1.4 | 1.33 | 1.33 | 0 | 0 | 0 | 0 | 1.4 | 1.33 | 1.33 | 1.4 | 1.33 | 1.33 |  |  | 48 |  |  |  |  | 1.4 | 1.33 | 1.33 |  |  |  |  | 1.4 | 1.33 | 1.33 |  |  |  |  | 1.4 | 1.33 | 1.33 |  |  |  |  | 1.4 | 1.33 | 1.33 | 1.4 | 1.33 | 1.33 |  |  | 48 |  |  |  |  | 1.4 | 1.33 | 1.33 |  |  |  |  | 1.4 | 1.33 | 1.33 |  |  |  |  | 1.4 | 1.33 | 1.33 |  |  |  |  | 1.4 | 1.33 | 1.33 | 1.4 | 1.33 | 1.33
48.001 | # | # | 0 | 0 | 0 | 0 | 1.4 | 1.37 | 1.37 | 0 | 0 | 0 | 0 | 1.4 | 1.37 | 1.37 | 0 | 0 | 0 | 0 | 1.4 | 1.37 | 1.37 | 0 | 0 | 0 | 0 | 1.4 | 1.37 | 1.37 | 1.4 | 1.37 | 1.37 |  |  | 48.001 |  |  |  |  | 1.4 | 1.37 | 1.37 |  |  |  |  | 1.4 | 1.37 | 1.37 |  |  |  |  | 1.4 | 1.37 | 1.37 |  |  |  |  | 1.4 | 1.37 | 1.37 | 1.4 | 1.37 | 1.37 |  |  | 48.001 |  |  |  |  | 1.47 | 1.37 | 1.37 |  |  |  |  | 1.47 | 1.37 | 1.37 |  |  |  |  | 1.47 | 1.37 | 1.37 |  |  |  |  | 1.47 | 1.37 | 1.37 | 1.4 | 1.37 | 1.37
50 | # | # | 0 | 0 | 0 | 0 | 1.4 | 1.37 | 1.37 | 0 | 0 | 0 | 0 | 1.4 | 1.37 | 1.37 | 0 | 0 | 0 | 0 | 1.4 | 1.37 | 1.37 | 0 | 0 | 0 | 0 | 1.4 | 1.37 | 1.37 | 1.4 | 1.37 | 1.37 |  |  | 50 |  |  |  |  | 1.4 | 1.37 | 1.37 |  |  |  |  | 1.4 | 1.37 | 1.37 |  |  |  |  | 1.4 | 1.37 | 1.37 |  |  |  |  | 1.4 | 1.37 | 1.37 | 1.4 | 1.37 | 1.37 |  |  | 50 |  |  |  |  | 1.47 | 1.37 | 1.37 |  |  |  |  | 1.47 | 1.37 | 1.37 |  |  |  |  | 1.47 | 1.37 | 1.37 |  |  |  |  | 1.47 | 1.37 | 1.37 | 1.4 | 1.37 | 1.37
50.001 | # | # | 0 | 0 | 0 | 0 | 0 | 1.37 | 1.37 | 0 | 0 | 0 | 0 | 0 | 1.37 | 1.37 | 0 | 0 | 0 | 0 | 0 | 1.37 | 1.37 | 0 | 0 | 0 | 0 | 0 | 1.37 | 1.37 | 1.68 | 1.37 | 1.37 |  |  | 50.001 | [5 empty] | 1.37 | 1.37 | [5 empty] | 1.37 | 1.37 | [5 empty] | 1.37 | 1.37 | [5 empty] | 1.37 | 1.37 | 1.68 | 1.37 | 1.37 |  |  | 50.001 | [5 empty] | 1.37 | 1.37 | [5 empty] | 1.37 | 1.37 | [5 empty] | 1.37 | 1.37 | [5 empty] | 1.37 | 1.37 | 1.4 | 1.37 | 1.37
53 | # | # | 0 | 0 | 0 | 0 | 0 | 1.37 | 1.37 | 0 | 0 | 0 | 0 | 0 | 1.37 | 1.37 | 0 | 0 | 0 | 0 | 0 | 1.37 | 1.37 | 0 | 0 | 0 | 0 | 0 | 1.37 | 1.37 | 1.68 | 1.37 | 1.37 |  |  | 53 | [5 empty] | 1.37 | 1.37 | [5 empty] | 1.37 | 1.37 | [5 empty] | 1.37 | 1.37 | [5 empty] | 1.37 | 1.37 | 1.68 | 1.37 | 1.37 |  |  | 53 | [5 empty] | 1.37 | 1.37 | [5 empty] | 1.37 | 1.37 | [5 empty] | 1.37 | 1.37 | [5 empty] | 1.37 | 1.37 | 1.4 | 1.37 | 1.37
54 | # | # | 0 | 0 | 0 | 0 | 0 | 1.37 | 1.37 | 0 | 0 | 0 | 0 | 0 | 1.37 | 1.37 | 0 | 0 | 0 | 0 | 0 | 1.37 | 1.37 | 0 | 0 | 0 | 0 | 0 | 1.37 | 1.37 | 0 | 1.37 | 1.37 |  |  | 54 | [5 empty] | 1.37 | 1.37 | [5 empty] | 1.37 | 1.37 | [5 empty] | 1.37 | 1.37 | [5 empty] | 1.37 | 1.37 |  | 1.37 | 1.37 |  |  | 54 | [5 empty] | 1.37 | 1.37 | [5 empty] | 1.37 | 1.37 | [5 empty] | 1.37 | 1.37 | [5 empty] | 1.37 | 1.37 |  | 1.37 | 1.37
54.001 | # | # | 0 | 0 | 0 | 0 | 0 | 0 | 1.41 | 0 | 0 | 0 | 0 | 0 | 0 | 1.41 | 0 | 0 | 0 | 0 | 0 | 0 | 1.41 | 0 | 0 | 0 | 0 | 0 | 0 | 1.41 | 0 | 0 | 1.41 |  |  | 54.001 | [6 empty] | 1.41 | [6 empty] | 1.41 | [6 empty] | 1.41 | [6 empty] | 1.41 |  |  | 1.41 |  |  | 54.001 | [6 empty] | 1.41 | [6 empty] | 1.41 | [6 empty] | 1.41 | [6 empty] | 1.41 |  |  | 1.41
60 | # | # | 0 | 0 | 0 | 0 | 0 | 0 | 1.41 | 0 | 0 | 0 | 0 | 0 | 0 | 1.41 | 0 | 0 | 0 | 0 | 0 | 0 | 1.41 | 0 | 0 | 0 | 0 | 0 | 0 | 1.41 | 0 | 0 | 1.41 |  |  | 60 | [6 empty] | 1.41 | [6 empty] | 1.41 | [6 empty] | 1.41 | [6 empty] | 1.41 |  |  | 1.41 |  |  | 60 | [6 empty] | 1.41 | [6 empty] | 1.41 | [6 empty] | 1.41 | [6 empty] | 1.41 |  |  | 1.41
60.001 | # | # | 0 | 0 | 0 | 0 | 0 | 0 | 1.63 | 0 | 0 | 0 | 0 | 0 | 0 | 1.63 | 0 | 0 | 0 | 0 | 0 | 0 | 1.63 | 0 | 0 | 0 | 0 | 0 | 0 | 1.63 | 0 | 0 | 1.63 |  |  | 60.001 | [6 empty] | 1.63 | [6 empty] | 1.63 | [6 empty] | 1.63 | [6 empty] | 1.63 |  |  | 1.63 |  |  | 60.001 | [6 empty] | 1.63 | [6 empty] | 1.63 | [6 empty] | 1.63 | [6 empty] | 1.63 |  |  | 1.63
67 | # | # | 0 | 0 | 0 | 0 | 0 | 0 | 1.63 | 0 | 0 | 0 | 0 | 0 | 0 | 1.63 | 0 | 0 | 0 | 0 | 0 | 0 | 1.63 | 0 | 0 | 0 | 0 | 0 | 0 | 1.63 | 0 | 0 | 1.63 |  |  | 67 | [6 empty] | 1.63 | [6 empty] | 1.63 | [6 empty] | 1.63 | [6 empty] | 1.63 |  |  | 1.63 |  |  | 67 | [6 empty] | 1.63 | [6 empty] | 1.63 | [6 empty] | 1.63 | [6 empty] | 1.63 |  |  | 1.63
67.001 | # | # | 0 | 0 | 0 | 0 | 0 | 0 | 2.23 | 0 | 0 | 0 | 0 | 0 | 0 | 2.23 | 0 | 0 | 0 | 0 | 0 | 0 | 2.23 | 0 | 0 | 0 | 0 | 0 | 0 | 2.23 | 0 | 0 | 2.23 |  |  | 67.001 | [6 empty] | 2.23 | [6 empty] | 2.23 | [6 empty] | 2.23 | [6 empty] | 2.23 |  |  | 2.23 |  |  | 67.001 | [6 empty] | 2.23 | [6 empty] | 2.23 | [6 empty] | 2.23 | [6 empty] | 2.23 |  |  | 2.23
70 | # | # | 0 | 0 | 0 | 0 | 0 | 0 | 2.23 | 0 | 0 | 0 | 0 | 0 | 0 | 2.23 | 0 | 0 | 0 | 0 | 0 | 0 | 2.23 | 0 | 0 | 0 | 0 | 0 | 0 | 2.23 | 0 | 0 | 2.23 |  |  | 70 | [6 empty] | 2.23 | [6 empty] | 2.23 | [6 empty] | 2.23 | [6 empty] | 2.23 |  |  | 2.23 |  |  | 70 | [6 empty] | 2.23 | [6 empty] | 2.23 | [6 empty] | 2.23 | [6 empty] | 2.23 |  |  | 2.23
| SWR | rho
Sensor max SWR | U2021XA: 1.22; U2022XA: 0.099099
DUT max SWR | U2021XA: 1.26; U2022XA: 0.11504
U2020 Average only mode (zero set, drift, noise calc) | BX58:AP58 U2057XA: U/L2050/60 Average only mode (zero set, drift, noise calc); BX58:AP58 L2054XA: U/L2050/60 Average only mode (zero set, drift, noise calc)
Zero set | U2021XA: 1e-08; BX58:AP58 U2057XA: Int zero set; BX58:AP58 U2061XA: 1e-09; BX58:AP58 U2063XA: Avg mode meas noise multiplier; BX58:AP58 L2054XA: Int zero set; BX58:AP58 L2055XA: 1e-09; BX58:AP58 L2057XA: Avg mode meas noise multiplier
Zero drift | U2021XA: 6e-09; BX58:AP58 U2057XA: Ext zero set (<300MHz); BX58:AP58 U2061XA: 9e-11; BX58:AP58 U2063XA: <=3.2s; BX58:AP58 U2064XA: =0.89*sqrt[1/(AvgCount*Aperture time)]; BX58:AP58 L2051XA: =; BX58:AP58 L2052XA: 3.9802; BX58:AP58 L2054XA: Ext zero set (<300MHz); BX58:AP58 L2055XA: 9e-11; BX58:AP58 L2057XA: <=3.2s; BX58:AP58 L2061XA: =0.89*sqrt[1/(AvgCount*Aperture time)]; U/L2051/52/53/61/62/63 (-70 to +15dBm) L2065XA: =; U/L2051/52/53/61/62/63 (-70 to +15dBm) L2066XA: 3.9802
Measurement noise | U2021XA: 4e-09; U2051XA: Noise multiplier =; U2053XA: 4.25; BX58:AP58 U2057XA: Ext zero set >=300MHz); BX58:AP58 U2061XA: 7e-11; BX58:AP58 U2063XA: >3.2s; BX58:AP58 U2064XA: =0.63*[1/(AvgCount*Aperture time)]^0.2; BX58:AP58 L2051XA: =; BX58:AP58 L2052XA: 1.147; BX58:AP58 L2054XA: Ext zero set >=300MHz); BX58:AP58 L2055XA: 7e-11; BX58:AP58 L2057XA: >3.2s; BX58:AP58 L2061XA: =0.63*[1/(AvgCount*Aperture time)]^0.2; U/L2051/52/53/61/62/63 (-70 to +15dBm) L2065XA: =; U/L2051/52/53/61/62/63 (-70 to +15dBm) L2066XA: 1.147
U2020 Measurement noise * noise multiplier | U2021XA: 1.7e-08; BX58:AP58 U2057XA: Zero drift; BX58:AP58 U2061XA: 2.5e-11; BX58:AP58 U2063XA: Noise_multiplier =; BX58:AP58 U2065XA: 3.9802; BX58:AP58 L2054XA: Zero drift; BX58:AP58 L2055XA: 2.5e-11; BX58:AP58 L2057XA: Noise_multiplier =; BX58:AP58 L2062XA: 3.9802
[9 empty] | Meas noise | 8e-11 |  | U/L2050/60 Avg mode noise = Meas_noise*Noise_Multiplier |  |  |  |  | = | 3.1842e-10 |  | Meas noise | 8e-11 |  | U/L2050/60 Avg mode noise = Meas_noise*Noise_Multiplier |  |  |  |  | = | 3.1842e-10
Avg Mode Meas Noise = | U2021XA: 1e-07; U2022XA: W; BX58:AP58 U2063XA: U/L2050/60 Avg mode noise / power; BX58:AP58 L2051XA: =; BX58:AP58 L2052XA: 3.184e-05%; BX58:AP58 L2057XA: U/L2050/60 Avg mode noise / power; U/L2051/52/53/61/62/63 (-70 to +15dBm) L2065XA: =; U/L2051/52/53/61/62/63 (-70 to +15dBm) L2066XA: 3.184e-05%
Zero drift = | U2021XA: 2.5e-11; U2022XA: W; BX58:AP58 U2063XA: (note: for avg mode, noise improvement is capped at 0.01%); BX58:AP58 L2051XA: =; BX58:AP58 L2052XA: 1e-07; BX58:AP58 L2057XA: (note: for avg mode, noise improvement is capped at 0.01%); U/L2051/52/53/61/62/63 (-70 to +15dBm) L2065XA: =; U/L2051/52/53/61/62/63 (-70 to +15dBm) L2066XA: 1e-07
Zero set = | U2021XA: 7e-11; U2022XA: W
Noise multiplier (average only mode): | BX58:AP58 U2063XA: Ext Zero set =; BX58:AP58 U2064XA: 7e-11; BX58:AP58 L2057XA: Ext Zero set =; BX58:AP58 L2061XA: 7e-11
Average setting | U2021XA: 1; U2022XA: 2; U2051XA: 4; U2052XA: 8; U2053XA: 16; BW58:AO58 U2054XA: 32; BX58:AP58 U2055XA: 64; BX58:AP58 U2056XA: 128; BX58:AP58 U2057XA: 256; BX58:AP58 U2061XA: 512; BX58:AP58 U2062XA: 1024; BX58:AP58 U2063XA: U2040 Zero set =; BX58:AP58 U2064XA: 7e-11; BX58:AP58 U2065XA: W; BX58:AP58 L2057XA: U2065XA Zero set =; BX58:AP58 L2061XA: 7e-11; BX58:AP58 L2062XA: W
Normal speed | U2021XA: 4.25; U2022XA: 2.84; U2051XA: 2.15; U2052XA: 1.52; U2053XA: 1; BW58:AO58 U2054XA: 0.78; BX58:AP58 U2055XA: 0.71; BX58:AP58 U2056XA: 0.52; BX58:AP58 U2057XA: 0.5; BX58:AP58 U2061XA: 0.47; BX58:AP58 U2062XA: 0.42
Double speed | U2021XA: 5.88; U2022XA: 4; U2051XA: 2.93; U2052XA: 1.89; U2053XA: 1.56; BW58:AO58 U2054XA: 1; BX58:AP58 U2055XA: 0.73; BX58:AP58 U2056XA: 0.55; BX58:AP58 U2057XA: 0.52; BX58:AP58 U2061XA: 0.48; BX58:AP58 U2062XA: 0.44
Freq range checking
Freq
U2021XA
U2022XA
U2051XA
U2052XA
U2053XA
U2054XA
U2055XA
U2056XA
U2057XA
U2061XA
U2062XA
U2063XA
U2064XA
U2065XA
U2066XA
U2067XA
L2051XA
L2052XA
L2053XA
L2054XA
L2055XA
L2056XA
L2057XA
L2061XA
L2062XA
L2063XA
L2064XA
L2065XA
L2066XA
L2067XA
L2065XT
L2066XT
L2067XT
max operating freq
Power range checking | U2021XA: Normal mode; U2022XA: Average only mode; U2051XA: 15; U2052XA: 10
U2021XA
U2022XA
U2051XA
U2052XA
U2053XA
U2054XA
U2055XA
U2056XA | BW58:AO58 U2054XA: 50
U2057XA | BW58:AO58 U2054XA: 50; BX58:AP58 U2056XA: 54
U2061XA
U2062XA
U2063XA
U2064XA
U2065XA
U2066XA | BW58:AO58 U2054XA: 50
U2067XA | BW58:AO58 U2054XA: 50; BX58:AP58 U2056XA: 54
L2051XA
L2052XA
L2053XA
L2054XA
L2055XA
L2056XA | BW58:AO58 U2054XA: 50
L2057XA | BW58:AO58 U2054XA: 50; BX58:AP58 U2056XA: 54
L2061XA
L2062XA
L2063XA
L2064XA
L2065XA
L2066XA | BW58:AO58 U2054XA: 50
L2067XA | BW58:AO58 U2054XA: 50; BX58:AP58 U2056XA: 54
L2065XT
L2066XT | BW58:AO58 U2054XA: 50
L2067XT | BW58:AO58 U2054XA: 50; BX58:AP58 U2056XA: 54
Relative Humidity > 70% | U2021XA: Yes; U2022XA: No
U2040 calibration uncertainty
Average mode | BX58:AP58 U2055XA: Normal mode; BX58:AP58 U2056XA: Video Off / High; BX58:AP58 U2061XA: (add 1.6% for RH 70 to 95%); BX58:AP58 U2064XA: Normal mode; BX58:AP58 U2065XA: Video Med/Low
| U2041XA | U2042XA | U2043XA | U2044XA | U2049XA |  | U2041XA | U2042XA | U2043XA | U2044XA | U2049XA |  | U2041XA | U2042XA | U2043XA | U2044XA | U2049XA
0.01 | U2021XA: 4.4%; U2022XA: 4.4%; U2051XA: 4.4%; U2052XA: 4.4%; U2053XA: 4.5%; BW58:AO58 U2054XA: 0.01; BX58:AP58 U2055XA: #; BX58:AP58 U2056XA: 5.7%; BX58:AP58 U2057XA: #; BX58:AP58 U2061XA: 5.7%; BX58:AP58 U2062XA: 4.5%; BX58:AP58 U2063XA: 0.01; BX58:AP58 U2064XA: #; BX58:AP58 U2065XA: 4.4%; BX58:AP58 U2066XA: #; BX58:AP58 U2067XA: 4.4%; BX58:AP58 L2051XA: 4.5%
0.03 | U2021XA: 4.4%; U2022XA: 4.4%; U2051XA: 4.4%; U2052XA: 4.4%; U2053XA: 4.5%; BW58:AO58 U2054XA: 0.03; BX58:AP58 U2055XA: #; BX58:AP58 U2056XA: 5.7%; BX58:AP58 U2057XA: #; BX58:AP58 U2061XA: 5.7%; BX58:AP58 U2062XA: 4.5%; BX58:AP58 U2063XA: 0.03; BX58:AP58 U2064XA: #; BX58:AP58 U2065XA: 4.4%; BX58:AP58 U2066XA: #; BX58:AP58 U2067XA: 4.4%; BX58:AP58 L2051XA: 4.5%
0.5 | U2021XA: 3.7%; U2022XA: 3.7%; U2051XA: 3.7%; U2052XA: 3.7%; U2053XA: 3.9%; BW58:AO58 U2054XA: 0.5; BX58:AP58 U2055XA: #; BX58:AP58 U2056XA: 5.2%; BX58:AP58 U2057XA: #; BX58:AP58 U2061XA: 5.2%; BX58:AP58 U2062XA: 4.1%; BX58:AP58 U2063XA: 0.5; BX58:AP58 U2064XA: #; BX58:AP58 U2065XA: 3.7%; BX58:AP58 U2066XA: #; BX58:AP58 U2067XA: 3.7%; BX58:AP58 L2051XA: 3.9%
1 | U2021XA: 3.7%; U2022XA: 3.7%; U2051XA: 3.7%; U2052XA: 3.7%; U2053XA: 3.8%; BW58:AO58 U2054XA: 1; BX58:AP58 U2055XA: #; BX58:AP58 U2056XA: 5.2%; BX58:AP58 U2057XA: #; BX58:AP58 U2061XA: 5.2%; BX58:AP58 U2062XA: 3.9%; BX58:AP58 U2063XA: 1; BX58:AP58 U2064XA: #; BX58:AP58 U2065XA: 3.7%; BX58:AP58 U2066XA: #; BX58:AP58 U2067XA: 3.7%; BX58:AP58 L2051XA: 3.9%
6 | U2021XA: 3.7%; U2022XA: 3.7%; U2051XA: 3.7%; U2052XA: 3.7%; U2053XA: 3.9%; BW58:AO58 U2054XA: 6; BX58:AP58 U2055XA: #; BX58:AP58 U2056XA: 5.3%; BX58:AP58 U2057XA: #; BX58:AP58 U2061XA: 5.3%; BX58:AP58 U2062XA: 4%; BX58:AP58 U2063XA: 6; BX58:AP58 U2064XA: #; BX58:AP58 U2065XA: 3.7%; BX58:AP58 U2066XA: #; BX58:AP58 U2067XA: 3.7%; BX58:AP58 L2051XA: 4%
10 | U2021XA: #; U2022XA: #; U2051XA: 3.7%; U2052XA: 3.7%; U2053XA: 4%; BW58:AO58 U2054XA: 10; BX58:AP58 U2055XA: #; BX58:AP58 U2056XA: #; BX58:AP58 U2057XA: #; BX58:AP58 U2061XA: 5.3%; BX58:AP58 U2062XA: 4.1%; BX58:AP58 U2063XA: 10; BX58:AP58 U2064XA: #; BX58:AP58 U2065XA: #; BX58:AP58 U2066XA: #; BX58:AP58 U2067XA: 3.7%; BX58:AP58 L2051XA: 4.1%
18 | U2021XA: #; U2022XA: #; U2051XA: 4%; U2052XA: 4%; U2053XA: 4.2%; BW58:AO58 U2054XA: 18; BX58:AP58 U2055XA: #; BX58:AP58 U2056XA: #; BX58:AP58 U2057XA: #; BX58:AP58 U2061XA: 5.4%; BX58:AP58 U2062XA: 4.3%; BX58:AP58 U2063XA: 18; BX58:AP58 U2064XA: #; BX58:AP58 U2065XA: #; BX58:AP58 U2066XA: #; BX58:AP58 U2067XA: 4%; BX58:AP58 L2051XA: 4.2%
26.5 | U2021XA: #; U2022XA: #; U2051XA: #; U2052XA: #; U2053XA: 4.9%; BW58:AO58 U2054XA: 26.5; BX58:AP58 U2055XA: #; BX58:AP58 U2056XA: #; BX58:AP58 U2057XA: #; BX58:AP58 U2061XA: #; BX58:AP58 U2062XA: 5%; BX58:AP58 U2063XA: 26.5; BX58:AP58 U2064XA: #; BX58:AP58 U2065XA: #; BX58:AP58 U2066XA: #; BX58:AP58 U2067XA: #; BX58:AP58 L2051XA: 4.9%
33 | U2021XA: #; U2022XA: #; U2051XA: #; U2052XA: #; U2053XA: 5.6%; BW58:AO58 U2054XA: 33; BX58:AP58 U2055XA: #; BX58:AP58 U2056XA: #; BX58:AP58 U2057XA: #; BX58:AP58 U2061XA: #; BX58:AP58 U2062XA: 5.7%; BX58:AP58 U2063XA: 33; BX58:AP58 U2064XA: #; BX58:AP58 U2065XA: #; BX58:AP58 U2066XA: #; BX58:AP58 U2067XA: #; BX58:AP58 L2051XA: 5.6%
Normal mode
| U2041XA | U2042XA | U2043XA | U2044XA | U2049XA |  | U2040 Cal Factor = |  | #N/A
0.01 | U2021XA: #REF!; U2022XA: #REF!; U2051XA: #REF!; U2052XA: #REF!; U2053XA: #REF!
0.03 | U2021XA: #REF!; U2022XA: #REF!; U2051XA: #REF!; U2052XA: #REF!; U2053XA: #REF!
0.5 | U2021XA: #REF!; U2022XA: #REF!; U2051XA: #REF!; U2052XA: #REF!; U2053XA: #REF!
1 | U2021XA: #REF!; U2022XA: #REF!; U2051XA: #REF!; U2052XA: #REF!; U2053XA: #REF!
6 | U2021XA: #REF!; U2022XA: #REF!; U2051XA: #REF!; U2052XA: #REF!; U2053XA: #REF!
10 | U2021XA: #REF!; U2022XA: #REF!; U2051XA: #REF!; U2052XA: #REF!; U2053XA: #REF!
18 | U2021XA: #REF!; U2022XA: #REF!; U2051XA: #REF!; U2052XA: #REF!; U2053XA: #REF!
26.5 | U2021XA: #REF!; U2022XA: #REF!; U2051XA: #REF!; U2052XA: #REF!; U2053XA: #REF!
33 | U2021XA: #REF!; U2022XA: #REF!; U2051XA: #REF!; U2052XA: #REF!; U2053XA: #REF!
U/L2050/60 calibration uncertainty
Average mode | [33 empty] | Normal mode (Video Off/High) | [18 empty] | Normal mode (Video Med/Low)
| U2051XA | U2052XA | U2053XA | U2054XA | U2055XA | U2056XA | U2057XA | U2061XA | U2062XA | U2063XA | U2064XA | U2065XA | U2066XA | U2067XA | L2051XA | L2052XA | L2053XA | L2054XA | L2055XA | L2056XA | L2057XA | L2061XA | L2062XA | L2063XA | L2064XA | L2065XA | L2066XA | L2067XA | L2065XT | L2066XT | L2067XT |  |  | U2061XA | U2062XA | U2063XA | U2064XA | L2065XA | U2066XA | L2067XA | L2061XA | L2062XA | L2063XA | L2064XA | L2065XA | L2066XA | L2067XA | L2065XT | L2066XT | L2067XT |  |  | U2061XA | U2062XA | U2063XA | U2064XA | U2065XA | U2066XA | U2067XA | L2061XA | L2062XA | L2063XA | L2064XA | L2065XA | L2066XA | L2067XA | L2065XT | L2066XT | L2067XT
0.01 | 4.3% | 4.3% | 4.4% | 4.6% | 4.6% | 4.8% | 4.8% | 4.3% | 4.3% | 4.4% | 4.6% | 4.6% | 4.8% | 4.8% | 4.3% | 4.3% | 4.4% | 4.6% | 4.6% | 4.8% | 4.8% | 4.3% | 4.3% | 4.4% | 4.6% | 4.6% | 4.8% | 4.8% | 4.7% | 4.8% | 4.8% |  | 0.01 | 4.3% | 4.3% | 4.4% | 4.7% | 4.7% | 5% | 5% | 4.3% | 4.3% | 4.4% | 4.7% | 4.7% | 5% | 5% | 4.6% | 5% | 5% |  | 0.01 | 4.5% | 4.5% | 4.3% | 4.4% | 4.4% | 4.5% | 4.5% | 4.5% | 4.5% | 4.3% | 4.4% | 4.4% | 4.5% | 4.5% | 4.7% | 4.5% | 4.5%
0.03 | 4.3% | 4.3% | 4.4% | 4.6% | 4.6% | 4.8% | 4.8% | 4.3% | 4.3% | 4.4% | 4.6% | 4.6% | 4.8% | 4.8% | 4.3% | 4.3% | 4.4% | 4.6% | 4.6% | 4.8% | 4.8% | 4.3% | 4.3% | 4.4% | 4.6% | 4.6% | 4.8% | 4.8% | 4.7% | 4.8% | 4.8% |  | 0.03 | 4.3% | 4.3% | 4.4% | 4.7% | 4.7% | 5% | 5% | 4.3% | 4.3% | 4.4% | 4.7% | 4.7% | 5% | 5% | 4.6% | 5% | 5% |  | 0.03 | 4.5% | 4.5% | 4.3% | 4.4% | 4.4% | 4.5% | 4.5% | 4.5% | 4.5% | 4.3% | 4.4% | 4.4% | 4.5% | 4.5% | 4.7% | 4.5% | 4.5%
0.031 | 3.5% | 3.5% | 3.9% | 3.6% | 3.6% | 3.2% | 3.2% | 3.5% | 3.5% | 3.9% | 3.6% | 3.6% | 3.2% | 3.2% | 3.5% | 3.5% | 3.9% | 3.6% | 3.6% | 3.2% | 3.2% | 3.5% | 3.5% | 3.9% | 3.6% | 3.6% | 3.2% | 3.2% | 3.8% | 3.2% | 3.2% |  | 0.031 | 3.6% | 3.6% | 4.1% | 4% | 4% | 4.2% | 4.2% | 3.6% | 3.6% | 4.1% | 4% | 4% | 4.2% | 4.2% | 4% | 4.2% | 4.2% |  | 0.031 | 3.8% | 3.8% | 4% | 3.5% | 3.5% | 3.7% | 3.7% | 3.8% | 3.8% | 4% | 3.5% | 3.5% | 3.7% | 3.7% | 3.9% | 3.7% | 3.7%
0.5 | 3.5% | 3.5% | 3.9% | 3.6% | 3.6% | 3.2% | 3.2% | 3.5% | 3.5% | 3.9% | 3.6% | 3.6% | 3.2% | 3.2% | 3.5% | 3.5% | 3.9% | 3.6% | 3.6% | 3.2% | 3.2% | 3.5% | 3.5% | 3.9% | 3.6% | 3.6% | 3.2% | 3.2% | 3.8% | 3.2% | 3.2% |  | 0.5 | 3.6% | 3.6% | 4.1% | 4% | 4% | 4.2% | 4.2% | 3.6% | 3.6% | 4.1% | 4% | 4% | 4.2% | 4.2% | 4% | 4.2% | 4.2% |  | 0.5 | 3.8% | 3.8% | 4% | 3.5% | 3.5% | 3.7% | 3.7% | 3.8% | 3.8% | 4% | 3.5% | 3.5% | 3.7% | 3.7% | 3.9% | 3.7% | 3.7%
0.501 | 3.5% | 3.5% | 3.9% | 3.6% | 3.6% | 3.6% | 3.6% | 3.5% | 3.5% | 3.9% | 3.6% | 3.6% | 3.6% | 3.6% | 3.5% | 3.5% | 3.9% | 3.6% | 3.6% | 3.6% | 3.6% | 3.5% | 3.5% | 3.9% | 3.6% | 3.6% | 3.6% | 3.6% | 3.9% | 3.6% | 3.6% |  | 0.501 | 3.6% | 3.6% | 4.1% | 4% | 4% | 4.2% | 4.2% | 3.6% | 3.6% | 4.1% | 4% | 4% | 4.2% | 4.2% | 4% | 4.2% | 4.2% |  | 0.501 | 3.8% | 3.8% | 4% | 3.5% | 3.5% | 3.9% | 3.9% | 3.8% | 3.8% | 4% | 3.5% | 3.5% | 3.9% | 3.9% | 4% | 3.9% | 3.9%
1 | 3.5% | 3.5% | 3.9% | 3.6% | 3.6% | 3.6% | 3.6% | 3.5% | 3.5% | 3.9% | 3.6% | 3.6% | 3.6% | 3.6% | 3.5% | 3.5% | 3.9% | 3.6% | 3.6% | 3.6% | 3.6% | 3.5% | 3.5% | 3.9% | 3.6% | 3.6% | 3.6% | 3.6% | 3.9% | 3.6% | 3.6% |  | 1 | 3.6% | 3.6% | 4.1% | 4% | 4% | 4.2% | 4.2% | 3.6% | 3.6% | 4.1% | 4% | 4% | 4.2% | 4.2% | 4% | 4.2% | 4.2% |  | 1 | 3.8% | 3.8% | 4% | 3.5% | 3.5% | 3.9% | 3.9% | 3.8% | 3.8% | 4% | 3.5% | 3.5% | 3.9% | 3.9% | 4% | 3.9% | 3.9%
1.001 | 3.5% | 3.5% | 3.9% | 3.6% | 3.6% | 3.6% | 3.6% | 3.5% | 3.5% | 3.9% | 3.6% | 3.6% | 3.6% | 3.6% | 3.5% | 3.5% | 3.9% | 3.6% | 3.6% | 3.6% | 3.6% | 3.5% | 3.5% | 3.9% | 3.6% | 3.6% | 3.6% | 3.6% | 3.9% | 3.6% | 3.6% |  | 1.001 | 3.6% | 3.6% | 4.1% | 4% | 4% | 4.2% | 4.2% | 3.6% | 3.6% | 4.1% | 4% | 4% | 4.2% | 4.2% | 4% | 4.2% | 4.2% |  | 1.001 | 3.7% | 3.7% | 4% | 3.5% | 3.5% | 3.9% | 3.9% | 3.7% | 3.7% | 4% | 3.5% | 3.5% | 3.9% | 3.9% | 4% | 3.9% | 3.9%
6 | 3.5% | 3.5% | 3.9% | 3.6% | 3.6% | 3.6% | 3.6% | 3.5% | 3.5% | 3.9% | 3.6% | 3.6% | 3.6% | 3.6% | 3.5% | 3.5% | 3.9% | 3.6% | 3.6% | 3.6% | 3.6% | 3.5% | 3.5% | 3.9% | 3.6% | 3.6% | 3.6% | 3.6% | 3.9% | 3.6% | 3.6% |  | 6 | 3.6% | 3.6% | 4.1% | 4% | 4% | 4.2% | 4.2% | 3.6% | 3.6% | 4.1% | 4% | 4% | 4.2% | 4.2% | 4% | 4.2% | 4.2% |  | 6 | 3.7% | 3.7% | 4% | 3.5% | 3.5% | 3.9% | 3.9% | 3.7% | 3.7% | 4% | 3.5% | 3.5% | 3.9% | 3.9% | 4% | 3.9% | 3.9%
6.001 |  | 3.6% | 4% | 3.7% | 3.7% | 3.6% | 3.6% |  | 3.6% | 4% | 3.7% | 3.7% | 3.6% | 3.6% |  | 3.6% | 4% | 3.7% | 3.7% | 3.6% | 3.6% |  | 3.6% | 4% | 3.7% | 3.7% | 3.6% | 3.6% | 3.9% | 3.6% | 3.6% |  | 6.001 |  | 3.6% | 4.1% | 4.1% | 4.1% | 4.3% | 4.3% |  | 3.6% | 4.1% | 4.1% | 4.1% | 4.3% | 4.3% | 4.2% | 4.3% | 4.3% |  | 6.001 |  | 3.7% | 4.1% | 3.7% | 3.7% | 4% | 4% |  | 3.7% | 4.1% | 3.7% | 3.7% | 4% | 4% | 4% | 4% | 4%
8 |  | 3.6% | 4% | 3.7% | 3.7% | 3.6% | 3.6% |  | 3.6% | 4% | 3.7% | 3.7% | 3.6% | 3.6% |  | 3.6% | 4% | 3.7% | 3.7% | 3.6% | 3.6% |  | 3.6% | 4% | 3.7% | 3.7% | 3.6% | 3.6% | 3.9% | 3.6% | 3.6% |  | 8 |  | 3.6% | 4.1% | 4.1% | 4.1% | 4.3% | 4.3% |  | 3.6% | 4.1% | 4.1% | 4.1% | 4.3% | 4.3% | 4.2% | 4.3% | 4.3% |  | 8 |  | 3.7% | 4.1% | 3.7% | 3.7% | 4% | 4% |  | 3.7% | 4.1% | 3.7% | 3.7% | 4% | 4% | 4% | 4% | 4%
8.001 |  | 3.6% | 4% | 3.7% | 3.7% | 3.6% | 3.6% |  | 3.6% | 4% | 3.7% | 3.7% | 3.6% | 3.6% |  | 3.6% | 4% | 3.7% | 3.7% | 3.6% | 3.6% |  | 3.6% | 4% | 3.7% | 3.7% | 3.6% | 3.6% | 3.9% | 3.6% | 3.6% |  | 8.001 |  | 3.6% | 4.1% | 4.1% | 4.1% | 4.3% | 4.3% |  | 3.6% | 4.1% | 4.1% | 4.1% | 4.3% | 4.3% | 4.2% | 4.3% | 4.3% |  | 8.001 |  | 3.7% | 4.1% | 3.7% | 3.7% | 4% | 4% |  | 3.7% | 4.1% | 3.7% | 3.7% | 4% | 4% | 4% | 4% | 4%
10 |  | 3.6% | 4% | 3.7% | 3.7% | 3.6% | 3.6% |  | 3.6% | 4% | 3.7% | 3.7% | 3.6% | 3.6% |  | 3.6% | 4% | 3.7% | 3.7% | 3.6% | 3.6% |  | 3.6% | 4% | 3.7% | 3.7% | 3.6% | 3.6% | 3.9% | 3.6% | 3.6% |  | 10 |  | 3.6% | 4.1% | 4.1% | 4.1% | 4.3% | 4.3% |  | 3.6% | 4.1% | 4.1% | 4.1% | 4.3% | 4.3% | 4.2% | 4.3% | 4.3% |  | 10 |  | 3.7% | 4.1% | 3.7% | 3.7% | 4% | 4% |  | 3.7% | 4.1% | 3.7% | 3.7% | 4% | 4% | 4% | 4% | 4%
10.001 |  | 3.7% | 4.2% | 3.7% | 3.7% | 3.6% | 3.6% |  | 3.7% | 4.2% | 3.7% | 3.7% | 3.6% | 3.6% |  | 3.7% | 4.2% | 3.7% | 3.7% | 3.6% | 3.6% |  | 3.7% | 4.2% | 3.7% | 3.7% | 3.6% | 3.6% | 3.9% | 3.6% | 3.6% |  | 10.001 |  | 3.8% | 4.3% | 4.1% | 4.1% | 4.3% | 4.3% |  | 3.8% | 4.3% | 4.1% | 4.1% | 4.3% | 4.3% | 4.2% | 4.3% | 4.3% |  | 10.001 |  | 3.8% | 4.3% | 3.7% | 3.7% | 4% | 4% |  | 3.8% | 4.3% | 3.7% | 3.7% | 4% | 4% | 4% | 4% | 4%
12 |  | 3.7% | 4.2% | 3.7% | 3.7% | 3.6% | 3.6% |  | 3.7% | 4.2% | 3.7% | 3.7% | 3.6% | 3.6% |  | 3.7% | 4.2% | 3.7% | 3.7% | 3.6% | 3.6% |  | 3.7% | 4.2% | 3.7% | 3.7% | 3.6% | 3.6% | 3.9% | 3.6% | 3.6% |  | 12 |  | 3.8% | 4.3% | 4.1% | 4.1% | 4.3% | 4.3% |  | 3.8% | 4.3% | 4.1% | 4.1% | 4.3% | 4.3% | 4.2% | 4.3% | 4.3% |  | 12 |  | 3.8% | 4.3% | 3.7% | 3.7% | 4% | 4% |  | 3.8% | 4.3% | 3.7% | 3.7% | 4% | 4% | 4% | 4% | 4%
12.001 |  | 3.7% | 4.2% | 3.9% | 3.9% | 3.8% | 3.8% |  | 3.7% | 4.2% | 3.9% | 3.9% | 3.8% | 3.8% |  | 3.7% | 4.2% | 3.9% | 3.9% | 3.8% | 3.8% |  | 3.7% | 4.2% | 3.9% | 3.9% | 3.8% | 3.8% | 3.9% | 3.8% | 3.8% |  | 12.001 |  | 3.8% | 4.3% | 4.2% | 4.2% | 4.6% | 4.6% |  | 3.8% | 4.3% | 4.2% | 4.2% | 4.6% | 4.6% | 4.2% | 4.6% | 4.6% |  | 12.001 |  | 3.8% | 4.3% | 3.8% | 3.8% | 4.4% | 4.4% |  | 3.8% | 4.3% | 3.8% | 3.8% | 4.4% | 4.4% | 4.4% | 4.4% | 4.4%
16 |  | 3.7% | 4.2% | 3.9% | 3.9% | 3.8% | 3.8% |  | 3.7% | 4.2% | 3.9% | 3.9% | 3.8% | 3.8% |  | 3.7% | 4.2% | 3.9% | 3.9% | 3.8% | 3.8% |  | 3.7% | 4.2% | 3.9% | 3.9% | 3.8% | 3.8% | 3.9% | 3.8% | 3.8% |  | 16 |  | 3.8% | 4.3% | 4.2% | 4.2% | 4.6% | 4.6% |  | 3.8% | 4.3% | 4.2% | 4.2% | 4.6% | 4.6% | 4.2% | 4.6% | 4.6% |  | 16 |  | 3.8% | 4.3% | 3.8% | 3.8% | 4.4% | 4.4% |  | 3.8% | 4.3% | 3.8% | 3.8% | 4.4% | 4.4% | 4.4% | 4.4% | 4.4%
16.001 |  | 3.7% | 4.2% | 4.2% | 4.2% | 4% | 4% |  | 3.7% | 4.2% | 4.2% | 4.2% | 4% | 4% |  | 3.7% | 4.2% | 4.2% | 4.2% | 4% | 4% |  | 3.7% | 4.2% | 4.2% | 4.2% | 4% | 4% | 4.3% | 4% | 4% |  | 16.001 |  | 3.8% | 4.3% | 4.6% | 4.6% | 4.8% | 4.8% |  | 3.8% | 4.3% | 4.6% | 4.6% | 4.8% | 4.8% | 4.5% | 4.8% | 4.8% |  | 16.001 |  | 3.8% | 4.3% | 4% | 4% | 4.7% | 4.7% |  | 3.8% | 4.3% | 4% | 4% | 4.7% | 4.7% | 5% | 4.7% | 4.7%
18 |  | 3.7% | 4.2% | 4.2% | 4.2% | 4% | 4% |  | 3.7% | 4.2% | 4.2% | 4.2% | 4% | 4% |  | 3.7% | 4.2% | 4.2% | 4.2% | 4% | 4% |  | 3.7% | 4.2% | 4.2% | 4.2% | 4% | 4% | 4.3% | 4% | 4% |  | 18 |  | 3.8% | 4.3% | 4.6% | 4.6% | 4.8% | 4.8% |  | 3.8% | 4.3% | 4.6% | 4.6% | 4.8% | 4.8% | 4.5% | 4.8% | 4.8% |  | 18 |  | 3.8% | 4.3% | 4% | 4% | 4.7% | 4.7% |  | 3.8% | 4.3% | 4% | 4% | 4.7% | 4.7% | 5% | 4.7% | 4.7%
18.001 |  |  | 4.5% | 4.2% | 4.2% | 4% | 4% |  |  | 4.5% | 4.2% | 4.2% | 4% | 4% |  |  | 4.5% | 4.2% | 4.2% | 4% | 4% |  |  | 4.5% | 4.2% | 4.2% | 4% | 4% | 4.3% | 4% | 4% |  | 18.001 |  |  | 4.6% | 4.6% | 4.6% | 4.8% | 4.8% |  |  | 4.6% | 4.6% | 4.6% | 4.8% | 4.8% | 4.5% | 4.8% | 4.8% |  | 18.001 |  |  | 4.5% | 4% | 4% | 4.7% | 4.7% |  |  | 4.5% | 4% | 4% | 4.7% | 4.7% | 5% | 4.7% | 4.7%
26.5 |  |  | 4.5% | 4.2% | 4.2% | 4% | 4% |  |  | 4.5% | 4.2% | 4.2% | 4% | 4% |  |  | 4.5% | 4.2% | 4.2% | 4% | 4% |  |  | 4.5% | 4.2% | 4.2% | 4% | 4% | 4.3% | 4% | 4% |  | 26.5 |  |  | 4.6% | 4.6% | 4.6% | 4.8% | 4.8% |  |  | 4.6% | 4.6% | 4.6% | 4.8% | 4.8% | 4.5% | 4.8% | 4.8% |  | 26.5 |  |  | 4.5% | 4% | 4% | 4.7% | 4.7% |  |  | 4.5% | 4% | 4% | 4.7% | 4.7% | 5% | 4.7% | 4.7%
26.501 |  |  | 5.1% | 4.3% | 4.3% | 4.5% | 4.5% |  |  | 5.1% | 4.3% | 4.3% | 4.5% | 4.5% |  |  | 5.1% | 4.3% | 4.3% | 4.5% | 4.5% |  |  | 5.1% | 4.3% | 4.3% | 4.5% | 4.5% | 4.9% | 4.5% | 4.5% |  | 26.501 |  |  | 5.2% | 4.7% | 4.7% | 5.1% | 5.1% |  |  | 5.2% | 4.7% | 4.7% | 5.1% | 5.1% | 5.1% | 5.1% | 5.1% |  | 26.501 |  |  | 5.2% | 4.2% | 4.2% | 4.8% | 4.8% |  |  | 5.2% | 4.2% | 4.2% | 4.8% | 4.8% | 5% | 4.8% | 4.8%
33 |  |  | 5.1% | 4.3% | 4.3% | 4.5% | 4.5% |  |  | 5.1% | 4.3% | 4.3% | 4.5% | 4.5% |  |  | 5.1% | 4.3% | 4.3% | 4.5% | 4.5% |  |  | 5.1% | 4.3% | 4.3% | 4.5% | 4.5% | 4.9% | 4.5% | 4.5% |  | 33 |  |  | 5.2% | 4.7% | 4.7% | 5.1% | 5.1% |  |  | 5.2% | 4.7% | 4.7% | 5.1% | 5.1% | 5.1% | 5.1% | 5.1% |  | 33 |  |  | 5.2% | 4.2% | 4.2% | 4.8% | 4.8% |  |  | 5.2% | 4.2% | 4.2% | 4.8% | 4.8% | 5% | 4.8% | 4.8%
33.001 |  |  |  | 4.8% | 4.8% | 4.7% | 4.7% |  |  |  | 4.8% | 4.8% | 4.7% | 4.7% |  |  |  | 4.8% | 4.8% | 4.7% | 4.7% |  |  |  | 4.8% | 4.8% | 4.7% | 4.7% | 5% | 4.7% | 4.7% |  | 33.001 |  |  |  | 5.3% | 5.3% | 5.1% | 5.1% |  |  |  | 5.3% | 5.3% | 5.1% | 5.1% | 5.1% | 5.1% | 5.1% |  | 33.001 |  |  |  | 4.7% | 4.7% | 4.8% | 4.8% |  |  |  | 4.7% | 4.7% | 4.8% | 4.8% | 5% | 4.8% | 4.8%
40 |  |  |  | 4.8% | 4.8% | 4.7% | 4.7% |  |  |  | 4.8% | 4.8% | 4.7% | 4.7% |  |  |  | 4.8% | 4.8% | 4.7% | 4.7% |  |  |  | 4.8% | 4.8% | 4.7% | 4.7% | 5% | 4.7% | 4.7% |  | 40 |  |  |  | 5.3% | 5.3% | 5.1% | 5.1% |  |  |  | 5.3% | 5.3% | 5.1% | 5.1% | 5.1% | 5.1% | 5.1% |  | 40 |  |  |  | 4.7% | 4.7% | 4.8% | 4.8% |  |  |  | 4.7% | 4.7% | 4.8% | 4.8% | 5% | 4.8% | 4.8%
40.001 | U2055XA: 5%; U2056XA: 5.1%; U2057XA: 5.1%; U2065XA: 5%; U2066XA: 5.1%; U2067XA: 5.1%; L2055XA: 5%; L2056XA: 5.1%; L2057XA: 5.1%; L2065XA: 5%; L2066XA: 5.1%; L2067XA: 5.1%; L2065XT: 5.6%; L2066XT: 5.1%; L2067XT: 5.1% |  | 40.001 |  |  |  |  | 5.7% | 5.8% | 5.8% |  |  |  |  | 5.7% | 5.8% | 5.8% | 5.8% | 5.8% | 5.8% |  | 40.001 |  |  |  |  | 4.9% | 5.2% | 5.2% |  |  |  |  | 4.9% | 5.2% | 5.2% | 5.6% | 5.2% | 5.2%
48 | U2055XA: 5%; U2056XA: 5.1%; U2057XA: 5.1%; U2065XA: 5%; U2066XA: 5.1%; U2067XA: 5.1%; L2055XA: 5%; L2056XA: 5.1%; L2057XA: 5.1%; L2065XA: 5%; L2066XA: 5.1%; L2067XA: 5.1%; L2065XT: 5.6%; L2066XT: 5.1%; L2067XT: 5.1% |  | 48 |  |  |  |  | 5.7% | 5.8% | 5.8% |  |  |  |  | 5.7% | 5.8% | 5.8% | 5.8% | 5.8% | 5.8% |  | 48 |  |  |  |  | 4.9% | 5.2% | 5.2% |  |  |  |  | 4.9% | 5.2% | 5.2% | 5.6% | 5.2% | 5.2%
48.001 | U2055XA: 5%; U2056XA: 5.1%; U2057XA: 5.1%; U2065XA: 5%; U2066XA: 5.1%; U2067XA: 5.1%; L2055XA: 5%; L2056XA: 5.1%; L2057XA: 5.1%; L2065XA: 5%; L2066XA: 5.1%; L2067XA: 5.1%; L2065XT: 5.6%; L2066XT: 5.1%; L2067XT: 5.1% |  | 48.001 |  |  |  |  | 5.7% | 5.9% | 5.9% |  |  |  |  | 5.7% | 5.9% | 5.9% | 5.8% | 5.9% | 5.9% |  | 48.001 |  |  |  |  | 4.9% | 5.2% | 5.2% |  |  |  |  | 4.9% | 5.2% | 5.2% | 5.6% | 5.2% | 5.2%
50 | U2055XA: 5%; U2056XA: 5.1%; U2057XA: 5.1%; U2065XA: 5%; U2066XA: 5.1%; U2067XA: 5.1%; L2055XA: 5%; L2056XA: 5.1%; L2057XA: 5.1%; L2065XA: 5%; L2066XA: 5.1%; L2067XA: 5.1%; L2065XT: 5.6%; L2066XT: 5.1%; L2067XT: 5.1% |  | 50 |  |  |  |  | 5.7% | 5.9% | 5.9% |  |  |  |  | 5.7% | 5.9% | 5.9% | 5.8% | 5.9% | 5.9% |  | 50 |  |  |  |  | 4.9% | 5.2% | 5.2% |  |  |  |  | 4.9% | 5.2% | 5.2% | 5.6% | 5.2% | 5.2%
50.001 | U2056XA: 5.1%; U2057XA: 5.1%; U2066XA: 5.1%; U2067XA: 5.1%; L2056XA: 5.1%; L2057XA: 5.1%; L2066XA: 5.1%; L2067XA: 5.1%; L2065XT: 5.8%; L2066XT: 5.1%; L2067XT: 5.1% |  | 50.001 | [5 empty] | 5.9% | 5.9% | [5 empty] | 5.9% | 5.9% | 5.9% | 5.9% | 5.9% |  | 50.001 | [5 empty] | 5.2% | 5.2% | [5 empty] | 5.2% | 5.2% | 5.8% | 5.2% | 5.2%
53 | U2056XA: 5.1%; U2057XA: 5.1%; U2066XA: 5.1%; U2067XA: 5.1%; L2056XA: 5.1%; L2057XA: 5.1%; L2066XA: 5.1%; L2067XA: 5.1%; L2065XT: 5.8%; L2066XT: 5.1%; L2067XT: 5.1% |  | 53 | [5 empty] | 5.9% | 5.9% | [5 empty] | 5.9% | 5.9% | 5.9% | 5.9% | 5.9% |  | 53 | [5 empty] | 5.2% | 5.2% | [5 empty] | 5.2% | 5.2% | 5.8% | 5.2% | 5.2%
54 | U2056XA: 5.1%; U2057XA: 5.1%; U2066XA: 5.1%; U2067XA: 5.1%; L2056XA: 5.1%; L2057XA: 5.1%; L2066XA: 5.1%; L2067XA: 5.1%; L2066XT: 5.1%; L2067XT: 5.1% |  | 54 | [5 empty] | 5.9% | 5.9% | [5 empty] | 5.9% | 5.9% |  | 5.9% | 5.9% |  | 54 | [5 empty] | 5.2% | 5.2% | [5 empty] | 5.2% | 5.2% |  | 5.2% | 5.2%
54.001 | U2057XA: 5.4%; U2067XA: 5.4%; L2057XA: 5.4%; L2067XA: 5.4%; L2067XT: 5.4% |  | 54.001 | [6 empty] | 6% | [6 empty] | 6% |  |  | 6% |  | 54.001 | [6 empty] | 5.3% | [6 empty] | 5.3% |  |  | 5.3%
60 | U2057XA: 5.4%; U2067XA: 5.4%; L2057XA: 5.4%; L2067XA: 5.4%; L2067XT: 5.4% |  | 60 | [6 empty] | 6% | [6 empty] | 6% |  |  | 6% |  | 60 | [6 empty] | 5.3% | [6 empty] | 5.3% |  |  | 5.3%
60.001 | U2057XA: 5.4%; U2067XA: 5.4%; L2057XA: 5.4%; L2067XA: 5.4%; L2067XT: 5.4% |  | 60.001 | [6 empty] | 6% | [6 empty] | 6% |  |  | 6% |  | 60.001 | [6 empty] | 5.3% | [6 empty] | 5.3% |  |  | 5.3%
67 | U2057XA: 5.4%; U2067XA: 5.4%; L2057XA: 5.4%; L2067XA: 5.4%; L2067XT: 5.4% |  | 67 | [6 empty] | 6% | [6 empty] | 6% |  |  | 6% |  | 67 | [6 empty] | 5.3% | [6 empty] | 5.3% |  |  | 5.3%
67.001 | U2057XA: 10.7%; U2067XA: 10.7%; L2057XA: 10.7%; L2067XA: 10.7%; L2067XT: 10.7% |  | 67.001 | [6 empty] | 11.2% | [6 empty] | 11.2% |  |  | 11.2% |  | 67.001 | [6 empty] | 10.8% | [6 empty] | 10.8% |  |  | 10.8%
70 | U2057XA: 10.7%; U2067XA: 10.7%; L2057XA: 10.7%; L2067XA: 10.7%; L2067XT: 10.7% |  | 70 | [6 empty] | 11.2% | [6 empty] | 11.2% |  |  | 11.2% |  | 70 | [6 empty] | 10.8% | [6 empty] | 10.8% |  |  | 10.8%
Normal mode | U2061XA: AI: BB; L2061XA: AP:BI
| U2051XA | U2052XA | U2053XA | U2054XA | U2055XA | U2056XA | U2057XA | U2061XA | U2062XA | U2063XA | U2064XA | U2065XA | U2066XA | U2067XA | L2051XA | L2052XA | L2053XA | L2054XA | L2055XA | L2056XA | L2057XA | L2061XA | L2062XA | L2063XA | L2064XA | L2065XA | L2066XA | L2067XA | L2065XT | L2066XT | L2067XT
0.01 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 4.3% | 4.3% | 4.4% | 4.7% | 4.7% | 5% | 5% | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 4.3% | 4.3% | 4.4% | 4.7% | 4.7% | 5% | 5% | 4.6% | 5% | 5% |  | U/L2050/60 Cal Factor = |  | 4.2%
0.03 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 4.3% | 4.3% | 4.4% | 4.7% | 4.7% | 5% | 5% | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 4.3% | 4.3% | 4.4% | 4.7% | 4.7% | 5% | 5% | 4.6% | 5% | 5%
0.031 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 3.6% | 3.6% | 4.1% | 4% | 4% | 4.2% | 4.2% | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 3.6% | 3.6% | 4.1% | 4% | 4% | 4.2% | 4.2% | 4% | 4.2% | 4.2%
0.5 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 3.6% | 3.6% | 4.1% | 4% | 4% | 4.2% | 4.2% | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 3.6% | 3.6% | 4.1% | 4% | 4% | 4.2% | 4.2% | 4% | 4.2% | 4.2%
0.501 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 3.6% | 3.6% | 4.1% | 4% | 4% | 4.2% | 4.2% | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 3.6% | 3.6% | 4.1% | 4% | 4% | 4.2% | 4.2% | 4% | 4.2% | 4.2%
1 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 3.6% | 3.6% | 4.1% | 4% | 4% | 4.2% | 4.2% | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 3.6% | 3.6% | 4.1% | 4% | 4% | 4.2% | 4.2% | 4% | 4.2% | 4.2%
1.001 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 3.6% | 3.6% | 4.1% | 4% | 4% | 4.2% | 4.2% | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 3.6% | 3.6% | 4.1% | 4% | 4% | 4.2% | 4.2% | 4% | 4.2% | 4.2%
6 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 3.6% | 3.6% | 4.1% | 4% | 4% | 4.2% | 4.2% | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 3.6% | 3.6% | 4.1% | 4% | 4% | 4.2% | 4.2% | 4% | 4.2% | 4.2% |  | Final |  | #N/A
6.001 | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | 3.6% | 4.1% | 4.1% | 4.1% | 4.3% | 4.3% | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | 3.6% | 4.1% | 4.1% | 4.1% | 4.3% | 4.3% | 4.2% | 4.3% | 4.3%
8 | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | 3.6% | 4.1% | 4.1% | 4.1% | 4.3% | 4.3% | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | 3.6% | 4.1% | 4.1% | 4.1% | 4.3% | 4.3% | 4.2% | 4.3% | 4.3%
8.001 | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | 3.6% | 4.1% | 4.1% | 4.1% | 4.3% | 4.3% | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | 3.6% | 4.1% | 4.1% | 4.1% | 4.3% | 4.3% | 4.2% | 4.3% | 4.3%
10 | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | 3.6% | 4.1% | 4.1% | 4.1% | 4.3% | 4.3% | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | 3.6% | 4.1% | 4.1% | 4.1% | 4.3% | 4.3% | 4.2% | 4.3% | 4.3%
10.001 | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | 3.8% | 4.3% | 4.1% | 4.1% | 4.3% | 4.3% | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | 3.8% | 4.3% | 4.1% | 4.1% | 4.3% | 4.3% | 4.2% | 4.3% | 4.3%
12 | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | 3.8% | 4.3% | 4.1% | 4.1% | 4.3% | 4.3% | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | 3.8% | 4.3% | 4.1% | 4.1% | 4.3% | 4.3% | 4.2% | 4.3% | 4.3%
12.001 | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | 3.8% | 4.3% | 4.2% | 4.2% | 4.6% | 4.6% | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | 3.8% | 4.3% | 4.2% | 4.2% | 4.6% | 4.6% | 4.2% | 4.6% | 4.6%
16 | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | 3.8% | 4.3% | 4.2% | 4.2% | 4.6% | 4.6% | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | 3.8% | 4.3% | 4.2% | 4.2% | 4.6% | 4.6% | 4.2% | 4.6% | 4.6%
16.001 | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | 3.8% | 4.3% | 4.6% | 4.6% | 4.8% | 4.8% | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | 3.8% | 4.3% | 4.6% | 4.6% | 4.8% | 4.8% | 4.5% | 4.8% | 4.8%
18 | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | 3.8% | 4.3% | 4.6% | 4.6% | 4.8% | 4.8% | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | 3.8% | 4.3% | 4.6% | 4.6% | 4.8% | 4.8% | 4.5% | 4.8% | 4.8%
18.001 | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  |  | 4.6% | 4.6% | 4.6% | 4.8% | 4.8% | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  |  | 4.6% | 4.6% | 4.6% | 4.8% | 4.8% | 4.5% | 4.8% | 4.8%
26.5 | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  |  | 4.6% | 4.6% | 4.6% | 4.8% | 4.8% | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  |  | 4.6% | 4.6% | 4.6% | 4.8% | 4.8% | 4.5% | 4.8% | 4.8%
26.501 | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  |  | 5.2% | 4.7% | 4.7% | 5.1% | 5.1% | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  |  | 5.2% | 4.7% | 4.7% | 5.1% | 5.1% | 5.1% | 5.1% | 5.1%
33 | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  |  | 5.2% | 4.7% | 4.7% | 5.1% | 5.1% | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  |  | 5.2% | 4.7% | 4.7% | 5.1% | 5.1% | 5.1% | 5.1% | 5.1%
33.001 | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  |  |  | 5.3% | 5.3% | 5.1% | 5.1% | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  |  |  | 5.3% | 5.3% | 5.1% | 5.1% | 5.1% | 5.1% | 5.1%
40 | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  |  |  | 5.3% | 5.3% | 5.1% | 5.1% | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  |  |  | 5.3% | 5.3% | 5.1% | 5.1% | 5.1% | 5.1% | 5.1%
40.001 | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  |  |  |  | 5.7% | 5.8% | 5.8% | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  |  |  |  | 5.7% | 5.8% | 5.8% | 5.8% | 5.8% | 5.8%
48 | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  |  |  |  | 5.7% | 5.8% | 5.8% | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  |  |  |  | 5.7% | 5.8% | 5.8% | 5.8% | 5.8% | 5.8%
48.001 | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  |  |  |  | 5.7% | 5.9% | 5.9% | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  |  |  |  | 5.7% | 5.9% | 5.9% | 5.8% | 5.9% | 5.9%
50 | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  |  |  |  | 5.7% | 5.9% | 5.9% | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  |  |  |  | 5.7% | 5.9% | 5.9% | 5.8% | 5.9% | 5.9%
50.001 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | [5 empty] | 5.9% | 5.9% | N/A | N/A | N/A | N/A | N/A | N/A | N/A | [5 empty] | 5.9% | 5.9% | 5.9% | 5.9% | 5.9%
53 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | [5 empty] | 5.9% | 5.9% | N/A | N/A | N/A | N/A | N/A | N/A | N/A | [5 empty] | 5.9% | 5.9% | 5.9% | 5.9% | 5.9%
54 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | [5 empty] | 5.9% | 5.9% | N/A | N/A | N/A | N/A | N/A | N/A | N/A | [5 empty] | 5.9% | 5.9% |  | 5.9% | 5.9%
54.001 | U2051XA: N/A; U2052XA: N/A; U2053XA: N/A; U2054XA: N/A; U2055XA: N/A; U2056XA: N/A; U2057XA: N/A; AI: BB U2067XA: 6%; AI: BB L2051XA: N/A; AI: BB L2052XA: N/A; AI: BB L2053XA: N/A; AI: BB L2054XA: N/A; AI: BB L2055XA: N/A; AI: BB L2056XA: N/A; AI: BB L2057XA: N/A; AP:BI L2067XA: 6%; AP:BI L2067XT: 6%
60 | U2051XA: N/A; U2052XA: N/A; U2053XA: N/A; U2054XA: N/A; U2055XA: N/A; U2056XA: N/A; U2057XA: N/A; AI: BB U2067XA: 6%; AI: BB L2051XA: N/A; AI: BB L2052XA: N/A; AI: BB L2053XA: N/A; AI: BB L2054XA: N/A; AI: BB L2055XA: N/A; AI: BB L2056XA: N/A; AI: BB L2057XA: N/A; AP:BI L2067XA: 6%; AP:BI L2067XT: 6%
60.001 | U2051XA: N/A; U2052XA: N/A; U2053XA: N/A; U2054XA: N/A; U2055XA: N/A; U2056XA: N/A; U2057XA: N/A; AI: BB U2067XA: 6%; AI: BB L2051XA: N/A; AI: BB L2052XA: N/A; AI: BB L2053XA: N/A; AI: BB L2054XA: N/A; AI: BB L2055XA: N/A; AI: BB L2056XA: N/A; AI: BB L2057XA: N/A; AP:BI L2067XA: 6%; AP:BI L2067XT: 6%
67 | U2051XA: N/A; U2052XA: N/A; U2053XA: N/A; U2054XA: N/A; U2055XA: N/A; U2056XA: N/A; U2057XA: N/A; AI: BB U2067XA: 6%; AI: BB L2051XA: N/A; AI: BB L2052XA: N/A; AI: BB L2053XA: N/A; AI: BB L2054XA: N/A; AI: BB L2055XA: N/A; AI: BB L2056XA: N/A; AI: BB L2057XA: N/A; AP:BI L2067XA: 6%; AP:BI L2067XT: 6%
67.001 | U2051XA: N/A; U2052XA: N/A; U2053XA: N/A; U2054XA: N/A; U2055XA: N/A; U2056XA: N/A; U2057XA: N/A; AI: BB U2067XA: 11.2%; AI: BB L2051XA: N/A; AI: BB L2052XA: N/A; AI: BB L2053XA: N/A; AI: BB L2054XA: N/A; AI: BB L2055XA: N/A; AI: BB L2056XA: N/A; AI: BB L2057XA: N/A; AP:BI L2067XA: 11.2%; AP:BI L2067XT: 11.2%
70 | U2051XA: N/A; U2052XA: N/A; U2053XA: N/A; U2054XA: N/A; U2055XA: N/A; U2056XA: N/A; U2057XA: N/A; AI: BB U2067XA: 11.2%; AI: BB L2051XA: N/A; AI: BB L2052XA: N/A; AI: BB L2053XA: N/A; AI: BB L2054XA: N/A; AI: BB L2055XA: N/A; AI: BB L2056XA: N/A; AI: BB L2057XA: N/A; AP:BI L2067XA: 11.2%; AP:BI L2067XT: 11.2%

## Revision history
Revision:
Rev 1 | 2010-10-01 | New creation for U/L2051/52/53/54/55/61/62/63/64/65 X-Series USB/LAN Peak and Average power sensors
Rev 2 | 2022-06-01 | Add new NPI of U/L2056/57XA, U/L2066/67XA & L2066/67XT X-Series USB/LAN Peak and Average power sensors
Rev 3 | 2022-09-12 | Remove error message