## Calculator
U2020 and U2040 X-Series USB Peak and Average Power Sensor
Uncertainty calculation for a power measurement (settled, average power)
Instruction: Fill up blue colour columns
1 | Power Level (dBm) | = | 0 |  | = | 0.001 | W
2 | Frequency (GHz) | = | 41
| Average setting | = | 1
| VBW setting | = | Off
| Free run / Trigger mode | = | Trigger mode
| Sensor mode | = | Average only mode | (Normal mode enables peak/average power measurements and trace view)
| Sensor model | = | U2022XA-H50
| DUT SWR | = | 1.26
| Gate length (us) - not required for Free Run or Average-only mode | = | 500
Fields applicable to U2040 X-Series only:
| (a) Aperture time (ms) | = | 50
| (b) Internal or External Zero | = | External Zero | (Internal zero - allow zeroing-on-the-go without switching off RF input power) (For U2040 X-Series, use External Zero when measure signal below -50dBm)
| (c) Relative Humidty > 70% | = | No
|  | Symbol | Value | Percentage
3 | Calculate sensor uncertainty due to measurement noise and zero drift
| Measurement noise | N | 1.7e-08 | 0.0017%
| Zero drift | D | 6e-09 | 0.0006%
| RSS of above terms | RSS(N,D) | -- | 0.001803%
4 | Zero uncertainty | Zs | 1e-08 | 0.001%
5 | Sensor calibration uncertainty | Kb | 7% | 7%
6 | System contribution, coverage factor of 2 | RSS(N, D, Zs, Kb) | -- | 7%
| (RSS of item 3, 4 and 5) | -- | --
7 | Standard uncertainty of mismatch | -- | --
| Sensor max SWR | ¦Гsensor¦ | 1.7
| DUT max SWR | ¦ГDUT¦ | 1.26
8 | Combined measurement uncertainty @ k = 1 | Uc | = | 4.086%
| k factor | k | = | 2
| Expanded Uncertainty | Uc * k | = | 8.173%
|  | Upper Limit Uncertainty | = | 0.34117
|  | Lower Limit Uncertainty | = | 0.37028
Revision 5
Date updated: 23 May 2024
© Keysight Technologies, Inc. 2012, 2014, 2015
The data used in this calculation is base on the specifications as in the U2020 & U2040 X-Series USB power sensors datasheet.
This calculation is base on ISO Guide to the Expression of Uncertainty in Measurement, often referred to as the GUM.
For more info on power meter and power sensor measurement uncertainty equation and calculation base on GUM, please refer to Keysight Application Note 1449-3 "Fundamental of RF and Microwave Power Measurements (Part 3), literature number 5988-9215EN.

## Data (hidden)
Measurement noise | Sensor mode
Free run mode | Normal mode
Trigger mode | Average only mode | [7 empty] | U2040 Normal mode
[10 empty] | Low/Med | High/off
Normal mode | [8 empty] | Meas noise | 1e-08 | 3.2e-08
Free run (U2040) | [8 empty] | Noise per sample | 1.5e-07 | 8e-07
Noise | 1e-07
Noise Multiplier | 1 | 2 | 4 | 8 | 16 | 32 | 64 | 128 | 256 | 512 | 1024
| 1 | 0.9 | 0.8 | 0.7 | 0.6 | 0.5 | 0.45 | 0.4 | 0.3 | 0.25 | 0.2
Free run measurement noise | = measurement noise x free run noise multiplier
(U2020) | 1e-07 | W | [6 empty] | Free run meas noise |  | =measurement noise x free run noise multiplier
Trigger mode | [8 empty] | (U2040) | = | 3.2e-08 | W
Noise per sample | 0.05 | 0.499 | 0.5 | 0.501 | 40
(U2020) | 3e-06 | 3e-06 | 2.5e-06 | 2.5e-06 | 2.5e-06
Noise per sample multiplier |  | <500MHz | >=500MHz
| Low | 0.6 | 0.55
| Med | 1.3 | 0.65
| High | 2.7 | 0.8
| Off | 1 | 1
Trigger mode measurement noise | = noise per sample x noise per sample multiplier x 1/sqrt[(gate length/12.5ns)] | [7 empty] | Trigger mode meas noise |  |  | = noise per sample x noise per saple multiplier x 1/sqrt[gate length/50ns]
(U2020) | 2.5e-06 | x | 1 | x | 0.005 |  |  |  | (U2040) |  | = | 8e-07 | x | 0.01
= | 1.25e-08 | W | [8 empty] | = | 8e-09 | W
Trigger mode measurement noise improvement limits at the measurement noise of 100 nW:
= | 1e-07
U2020 NORM Measurement noise = | 1e-07 | W | [6 empty] | U2040 NORM meas noise |  | = | 8e-09 | W | 5.6569e-08
NORM measurement noise = | 8e-09 | W
U2020 Final Measurement noise in % | = Noise / Power, noise power is capped for power > 100 uW, for these cases, use: Noise / 100 uW
= | 0.1% | [6 empty] | U2040 final NORM mode meas noise in % = |  |  |  | Noise / power
[11 empty] | = | 0.0008%
Final NORM meas noise in % | 0.0008%
NORM mode zero drift | U2021XA | U2022XA | U2041XA | U2042XA | U2043XA | U2044XA | U2049XA | U2022XA-H50
| 1e-07 | 2e-07 | 1.5e-08 | 1.5e-08 | 1.5e-08 | 1.5e-08 | 1.5e-08 | 2e-07
NORM mode zero drift = | 2e-07 | W
U2040 NORM Zero Set
NORM mode Zero set | U2021XA | U2022XA | U2041XA | U2042XA | U2043XA | U2044XA | U2049XA | U2022XA-H50 |  | External Zero | Internal Zero
| 2e-07 | 2e-07 | 5e-08 | 5e-08 | 5e-08 | 5e-08 | 5e-08 | 2e-07 | Low | 1.6e-08 | 2.3e-08
NORM mode zero set = | 2e-07 | W | [6 empty] | Med | 1.6e-08 | 2.3e-08
[9 empty] | High | 5e-08 | 6e-08
[9 empty] | Off | 5e-08 | 6e-08
Calibration uncertainty
| U2021XA | U2022XA | U2022XA-H50
0.05 | 4.2% | 4.3% | 4.3%
0.5 | 4.2% | 4.3% | 4.3%
0.5001 | 4% | 4.2% | 4.2%
1.001 | 4% | 4.5% | 4.5%
10.001 | 4.5% | 4.5% | 4.5%
18.001 | # | 5.3% | 5.3% | U2020 Cal Factor = |  | 7%
26.501 | # | 5.8% | 5.8%
40.001 | # | # | 7%
47.001 | # | # | 8%
50 | # | # | 8%
Sensor max SWR | [8 empty] | U2040 (-70 to +15dBm) | [6 empty] | U2040 (+15 to +26dBm)
| U2021XA | U2022XA | U2041XA | U2042XA | U2043XA | U2044XA | U2049XA | U2022XA-H50 |  | U2041XA | U2042XA | U2043XA | U2044XA | U2049XA |  |  | U2041XA | U2042XA | U2043XA | U2044XA | U2049XA
0.01 | # | # | 1.2 | 1.2 | 1.2 | 1.2 | 2.18 | # | 0.01 | 1.2 | 1.2 | 1.2 | 1.2 | 2.18 |  | 0.01 | 1.29 | 1.29 | 1.29 | 1.29 | 2.21
0.0301 | # | # | 1.2 | 1.2 | 1.2 | 1.2 | 1.35 | # | 0.0301 | 1.2 | 1.2 | 1.2 | 1.2 | 1.35 |  | 0.0301 | 1.29 | 1.29 | 1.29 | 1.29 | 1.37
0.05 | 1.2 | 1.2 | 1.2 | 1.2 | 1.2 | 1.2 | 1.35 | 1.2 | 0.05 | 1.2 | 1.2 | 1.2 | 1.2 | 1.35 |  | 0.05 | 1.29 | 1.29 | 1.29 | 1.29 | 1.37
0.0501 | 1.2 | 1.2 | 1.2 | 1.2 | 1.2 | 1.2 | 1.22 | 1.2 | 0.0501 | 1.2 | 1.2 | 1.2 | 1.2 | 1.22 |  | 0.0501 | 1.29 | 1.29 | 1.29 | 1.29 | 1.24
0.101 | 1.2 | 1.2 | 1.2 | 1.2 | 1.2 | 1.2 | 1.17 | 1.2 | 0.101 | 1.2 | 1.2 | 1.2 | 1.2 | 1.17 |  | 0.101 | 1.29 | 1.29 | 1.29 | 1.29 | 1.21
6 | 1.2 | 1.2 | 1.2 | 1.2 | 1.2 | 1.2 | 1.17 | 1.2 | 6 | 1.2 | 1.2 | 1.2 | 1.2 | 1.17 |  | 6 | 1.29 | 1.29 | 1.29 | 1.29 | 1.21
6.001 | 1.2 | 1.2 | # | # | 1.26 | 1.26 | 1.17 | 1.2 | 6.001 | # | # | 1.26 | 1.26 | 1.17 |  | 6.001 | # | # | 1.3 | 1.3 | 1.21
10 | 1.2 | 1.2 | # | # | 1.26 | 1.26 | 1.17 | 1.2 | 10 | # | # | 1.26 | 1.26 | 1.17 |  | 10 | # | # | 1.3 | 1.3 | 1.21
10.001 | 1.26 | 1.26 | # | # | 1.26 | 1.26 | 1.17 | 1.26 | 10.001 | # | # | 1.26 | 1.26 | 1.17 |  | 10.001 | # | # | 1.3 | 1.3 | 1.21
11.501 | 1.26 | 1.26 | # | # | 1.26 | 1.26 | 1.29 | 1.26 | 11.501 | # | # | 1.26 | 1.26 | 1.29 |  | 11.501 | # | # | 1.3 | 1.3 | 1.33
18 | 1.26 | 1.26 | # | # | 1.26 | 1.26 | 1.29 | 1.26 | 18 | # | # | 1.26 | 1.26 | 1.29 |  | 18 | # | # | 1.3 | 1.3 | 1.33
18.001 | # | 1.3 | # | # | # | # | 1.29 | 1.3 | 18.001 | # | # | # | # | 1.29 |  | 18.001 | # | # | # | # | 1.33
26.5 | # | 1.3 | # | # | # | # | 1.29 | 1.3 | 26.5 | # | # | # | # | 1.29 |  | 26.5 | # | # | # | # | 1.33
26.501 | # | 1.5 | # | # | # | # | 1.29 | 1.5 | 26.501 | # | # | # | # | 1.29 |  | 26.501 | # | # | # | # | 1.33
30.001 | # | 1.5 | # | # | # | # | 1.33 | 1.5 | 30.001 | # | # | # | # | 1.33 |  | 30.001 | # | # | # | # | 1.36
33 | # | 1.5 | # | # | # | # | 1.33 | 1.5 | 33 | # | # | # | # | 1.33 |  | 33 | # | # | # | # | 1.36
40 | # | 1.5 | # | # | # | # | # | 1.5 | 40 | # | # | # | # | # |  | 40 | # | # | # | # | #
40.001 | # | # | # | # | # | # | # | 1.7
50 | # | # | # | # | # | # | # | 1.7
| SWR | rho
Sensor max SWR | 1.7 | 0.25926
DUT max SWR | 1.26 | 0.11504
U2020 Average only mode (zero set, drift, noise calc) | [8 empty] | U2040 Average only mode (zero set, drift, noise calc)
Zero set | 1e-08 | [7 empty] | Int zero set | 1e-09 |  | Avg mode meas noise multiplier
Zero drift | 6e-09 | [7 empty] | Ext zero set (<300MHz) | 1e-10 |  | <=3.2s | =0.89*sqrt[1/(AvgCount*Aperture time)] |  |  |  | = | 3.9802
Measurement noise | 4e-09 |  | Noise multiplier = |  | 4.25 |  |  |  | Ext zero set >=300MHz) | 7e-11 |  | >3.2s | =0.63*[1/(AvgCount*Aperture time)]^0.2 |  |  |  | = | 1.147
U2020 Measurement noise * noise multiplier | 1.7e-08 | [7 empty] | Zero drift | 2.5e-11 |  | Noise_multiplier = |  | 3.9802
[9 empty] | Meas noise | 8e-11 |  | U2040 Avg mode noise = Meas_noise*Noise_Multiplier |  |  |  |  | = | 3.1842e-10
Avg Mode Meas Noise = | 1.7e-08 | W | [9 empty] | U2040 Avg mode noise / power |  |  |  |  | = | 3.184e-05%
Zero drift = | 6e-09 | W | [9 empty] | (note: for avg mode, noise improvement is capped at 0.01%) |  |  |  |  | = | 1e-07
Zero set = | 1e-08 | W
Noise multiplier (average only mode): | [11 empty] | Ext Zero set = | 7e-11
Average setting | 1 | 2 | 4 | 8 | 16 | 32 | 64 | 128 | 256 | 512 | 1024 | U2040 Zero set = | 7e-11 | W
Normal speed | 4.25 | 2.84 | 2.15 | 1.52 | 1 | 0.78 | 0.71 | 0.52 | 0.5 | 0.47 | 0.42
Double speed | 5.88 | 4 | 2.93 | 1.89 | 1.56 | 1 | 0.73 | 0.55 | 0.52 | 0.48 | 0.44
Freq range checking | Freq
U2021XA | Invalid input!
U2022XA | Invalid input!
U2041XA | Invalid input!
U2042XA | Invalid input!
U2043XA | Invalid input!
U2044XA | Invalid input!
U2049XA | Invalid input! |  | optio H50 = | * -45 dBm to 8 dBm (> 40 GHz to 50 GHz)
U2022XA-H50
Power range checking | Normal mode | Average only mode
U2021XA
U2022XA
U2041XA
U2042XA
U2043XA
U2044XA
U2049XA
U2022XA-H50
|  |  | normal mode chk | avg mode chk | pwr level chk
Relative Humidity > 70% | Yes | No
U2040 calibration uncertainty
Average mode | [6 empty] | Normal mode | Video Off / High |  | (add 1.6% for RH 70 to 95%) |  |  | Normal mode | Video Med/Low
| U2041XA | U2042XA | U2043XA | U2044XA | U2049XA |  | U2041XA | U2042XA | U2043XA | U2044XA | U2049XA |  | U2041XA | U2042XA | U2043XA | U2044XA | U2049XA
0.01 | 4.4% | 4.4% | 4.4% | 4.4% | 4.5% | 0.01 | # | 5.7% | # | 5.7% | 4.5% | 0.01 | # | 4.4% | # | 4.4% | 4.5%
0.03 | 4.4% | 4.4% | 4.4% | 4.4% | 4.5% | 0.03 | # | 5.7% | # | 5.7% | 4.5% | 0.03 | # | 4.4% | # | 4.4% | 4.5%
0.5 | 3.7% | 3.7% | 3.7% | 3.7% | 3.9% | 0.5 | # | 5.2% | # | 5.2% | 4.1% | 0.5 | # | 3.7% | # | 3.7% | 3.9%
1 | 3.7% | 3.7% | 3.7% | 3.7% | 3.8% | 1 | # | 5.2% | # | 5.2% | 3.9% | 1 | # | 3.7% | # | 3.7% | 3.9%
6 | 3.7% | 3.7% | 3.7% | 3.7% | 3.9% | 6 | # | 5.3% | # | 5.3% | 4% | 6 | # | 3.7% | # | 3.7% | 4%
10 | # | # | 3.7% | 3.7% | 4% | 10 | # | # | # | 5.3% | 4.1% | 10 | # | # | # | 3.7% | 4.1%
18 | # | # | 4% | 4% | 4.2% | 18 | # | # | # | 5.4% | 4.3% | 18 | # | # | # | 4% | 4.2%
26.5 | # | # | # | # | 4.9% | 26.5 | # | # | # | # | 5% | 26.5 | # | # | # | # | 4.9%
33 | # | # | # | # | 5.6% | 33 | # | # | # | # | 5.7% | 33 | # | # | # | # | 5.6%
Normal mode
| U2041XA | U2042XA | U2043XA | U2044XA | U2049XA |  | U2040 Cal Factor = |  | #N/A
0.01 | # | 5.7% | # | 5.7% | 4.5%
0.03 | # | 5.7% | # | 5.7% | 4.5%
0.5 | # | 5.2% | # | 5.2% | 4.1%
1 | # | 5.2% | # | 5.2% | 3.9%
6 | # | 5.3% | # | 5.3% | 4%
10 | # | # | # | 5.3% | 4.1%
18 | # | # | # | 5.4% | 4.3%
26.5 | # | # | # | # | 5%
33 | # | # | # | # | 5.7%

## Revision history
Revision:
Rev 1 | 2012-10-22 | New creation for U2020 X-Series
Rev 2 | 2014-02-07 | Added 'average only mode' with extended power range down to -45dBm
Rev 3 | 2015-11-25 | Added in U2040 X-Series
Rev 4 | 2017-09-07 | Fixed "dot comma" bug due to regional settings
Rev 5 | 2024-05-23 | To include U2022XA-H50 that freq up to 50GHz in MU calculation